## Documents
Slug | Name | Short Code | Sections | Requirements | Created
architecture-decisions | Architecture Decisions | ARC | 1 | 7 | 2026-03-30T22:53:38.344Z
interface-requirements | Interface Requirements | IFC | 9 | 31 | 2026-03-30T22:53:38.026Z
stakeholder-requirements | Stakeholder Requirements | STK | 1 | 18 | 2026-03-30T22:53:37.298Z
subsystem-requirements | Subsystem Requirements | SUB | 9 | 93 | 2026-03-30T22:53:38.187Z
system-requirements | System Requirements | SYS | 1 | 23 | 2026-03-30T22:53:37.822Z
verification-requirements | Verification Requirements | VER | 9 | 116 | 2026-03-30T22:53:38.502Z

## Requirements
Ref | Text | Document | Section | Pattern | Verification | QA Score | Compliance | Rationale | Tags
ARC-REQ-001 | The BCI Processing Subsystem SHALL implement combined EEG signal acquisition and intent classification on a single embedded processor, sharing the BLE headset interface and neural signal pipeline to achieve end-to-end latency <50ms. Alternative: separate acquisition and classification boards — rejected because the 50ms latency budget leaves insufficient margin for inter-board data transfer and additional hardware increases failure points. | architecture-decisions | Architecture Decisions |  | Analysis | 71 |  | Co-location of acquisition and classification on a single processor eliminates a high-bandwidth inter-subsystem link for raw 250Hz 32-channel EEG data. The 50ms end-to-end latency threshold is derived from BCI usability research (mean motor imagery detection lag must be <50ms for natural-feeling control). Separate boards would require ≥10ms DMA transfer overhead, consuming the entire latency margin. | architecture, bci, session-696, idempotency:arc-bci-696, informational
ARC-REQ-002 | ARC: Safety Subsystem — Independent safety channel with dedicated processor. IEC 61508 (Functional safety of E/E/PE safety-related systems) requires that SIL 3 safety functions be independent of the control path. The Safety Subsystem uses a separate watchdog processor and hardware-level motor power disconnect that operates even if the main application processor fails. Alternative considered: software-only safety monitor on the main processor — rejected because a single processor fault would disable both control and safety, violating SIL 3 independence. | architecture-decisions | Architecture Decisions |  | Inspection | 57 |  | IEC 61508 (Functional safety of E/E/PE safety-related systems) clause 7.4.2 requires that SIL 3 safety functions be implemented on hardware independent from the control path. A separate watchdog processor satisfies this independence requirement. Verifiable by inspection of the hardware architecture and IEC 61508 compliance matrix confirming independent power supply, independent clock, and independent activation path for motor power disconnect. | architecture, safety, session-696, idempotency:arc-safety-696, informational
ARC-REQ-003 | The Perception Subsystem SHALL use dedicated sensor hardware and processing independent of the Drive Subsystem, so that obstacle detection capability is preserved when the Drive Subsystem is in a fault state and Drive Subsystem operation is preserved when the Perception Subsystem fails. Alternative: integrated drive-and-perception controller — rejected because sensor processor lockup would simultaneously disable obstacle detection and emergency braking. | architecture-decisions | Architecture Decisions |  | Analysis | 86 |  | Failure independence between perception and drive is required under BS EN ISO 13482 (Safety requirements for personal care robots) clause 5.4 (hazardous failure avoidance). If a single processor controlled both functions, a software fault could disable the obstacle detection while the drive motors remain energised. Separation ensures the Safety Subsystem watchdog can still trigger E-stop via hardware timeout even with Perception failed. | architecture, perception, session-696, idempotency:arc-perception-696, informational
ARC-REQ-004 | ARC: Communication Subsystem — Separated from HMI to isolate wireless reliability from user feedback. BLE and Wi-Fi radios share antenna resources and protocol stacks but serve different latency requirements (500ms emergency alerts vs. background CMMS telemetry). Grouping all external data links together allows a single security boundary for TLS and GDPR compliance. Alternative considered: distributing BLE to BCI Processing and Wi-Fi to HMI — rejected because scattered radio hardware complicates EMC certification and security auditing. | architecture-decisions | Architecture Decisions |  | Inspection | 57 |  | BLE emergency alert latency (500ms SYS-REQ-017) and Wi-Fi CMMS telemetry (background, best-effort) have incompatible QoS requirements. Grouping both under Communication Subsystem allows a unified RF compliance boundary for CE marking and a single TLS/GDPR security review scope. Distributing radios across subsystems would require separate EMC test campaigns per subsystem and complicate GDPR data flow documentation. | architecture, communication, session-696, idempotency:arc-comms-696, informational
ARC-REQ-005 | The Power Subsystem architecture SHALL implement dual-authority thermal protection: the Battery Management System SHALL signal the Safety Subsystem when cell temperature exceeds 60°C or cell voltage imbalance exceeds 100mV, and the Safety Subsystem SHALL independently command the 48V disconnect relay. Alternative: BMS-only disconnect — rejected because a BMS processor fault would prevent disconnect during thermal runaway (H-003, catastrophic severity). | architecture-decisions | Architecture Decisions |  | Test | 86 |  | Dual-authority thermal protection is required because H-003 (battery thermal runaway) is classified as catastrophic under BS EN 62133-2 (Safety requirements for portable sealed secondary lithium cells). A single BMS processor failure must not prevent safe battery disconnect. IEC 61508 SIL 3 allocation for H-003 requires that no single-point failure in the safety function is credible — hence the Safety Subsystem holds independent temperature sensing and direct relay control. | architecture, power, safety, session-696, idempotency:arc-power-696, informational
ARC-REQ-006 | ARC: Drive Subsystem — Differential-drive 48V BLDC architecture with dedicated Motor Controller Unit. Four-component decomposition (Motor Controller Unit, Left/Right Drive Motor Assemblies, Drive Power Stage) chosen over a pre-integrated motor controller module. The dual-channel MCU architecture allows independent velocity closed-loop control on each motor, enabling differential steering without a mechanical differential. The Drive Power Stage is isolated from the 48V bus via the hardware Motor Power Isolation Relay (Safety Subsystem) rather than software-only interlock — this ensures safe-state is reachable even with MCU firmware failure, consistent with SIL 2 (IEC 61508) allocation. | architecture-decisions | Architecture Decisions |  | Inspection | 57 |  | SYS-REQ-003 requires speed limiting at 6km/h (Normal) and 2km/h (Restricted), which must be enforced in the Motor Controller Unit firmware. SYS-REQ-002 requires 250ms safe stop — hardware relay isolation ensures this is achievable independent of MCU state. Differential drive was selected over Ackermann steering because the buggy's narrow indoor operating environment requires a turning radius of <800mm, achievable only with zero-radius turns on differential drive. | architecture, drive-subsystem, session-700, idempotency:arc-drive-subsystem-700
ARC-REQ-007 | The Vehicle Platform architecture SHALL implement the Electronics Bay as a separate sealed IP54 enclosure mounted to the Chassis Frame, containing the CAN bus harness within a defined EMC boundary and enabling field serviceability without chassis disassembly. Alternative: electronics integrated into chassis weldment — rejected because weld-bonded enclosures prevent field serviceability and make EMC sealing impractical. | architecture-decisions | Architecture Decisions |  | Inspection | 71 |  | A separate IP54 Electronics Bay satisfies IEC 60529 ingress protection and MHRA MDR Annex I Essential Requirements for implant-adjacent electronics. Field serviceability is an operational requirement (Carer Maintenance scenario): technicians must be able to replace individual modules without returning the vehicle to the manufacturer. Integrating electronics into the chassis weldment would require certified welding inspection on every repair — disproportionate to the service risk. | architecture, vehicle-platform, session-701, idempotency:arc-vehicle-platform-701
IFC-REQ-005 | The interface between the EEG Accessibility Buggy and the Service Diagnostic Laptop SHALL use USB-C wired connection for firmware updates, diagnostic log export, configuration changes, and calibration data, protected by service authentication key with role-based access control. | interface-requirements | Interface Requirements |  | Demonstration | 86 |  | External interface: wired USB-C ensures reliable high-bandwidth data transfer for firmware images (dual-bank flash) and diagnostic logs. Authentication key prevents unauthorised firmware modification — critical for IEC 62304 software change control and preventing malicious firmware injection. | interface, external, maintenance, session-696, idempotency:ifc-ext-diagnostic-696
IFC-BCIPROCESSINGSUBSYSTEM-001 | The interface between the EEG Accessibility Buggy and the EEG Headset SHALL use Bluetooth Low Energy 5.0 to stream 16-32 channels of 24-bit EEG data at 250 Hz with end-to-end latency not exceeding 50ms from electrode to onboard processor. | interface-requirements | BCI Processing Subsystem |  | Test | 86 |  | External interface: the headset is the sole BCI input device. BLE 5.0 provides sufficient bandwidth (~2 Mbps) for 32×24bit×250Hz = 192 kbps with margin. 50ms latency budget is derived from the 200ms E-stop requirement minus processing and actuation time. | interface, external, bci, session-696, idempotency:ifc-ext-headset-696
IFC-BCIPROCESSINGSUBSYSTEM-002 | The interface between the EEG Acquisition Module and the Artifact Rejection Engine SHALL transfer 32-channel EEG sample arrays at 256 samples/second with a maximum inter-process latency of 5ms and sample loss rate not exceeding 0.1%. | interface-requirements | BCI Processing Subsystem |  | Test | 100 |  | 256 Hz is the headset sampling rate; all 32 channels required for spatial filtering in artifact rejection. 5ms latency budget is allocated from the 20ms artifact rejection window. 0.1% sample loss is the threshold at which artifact rejection algorithms begin to produce corrupted output windows, as gap-filling interpolation errors compound across the 1-second epoch. | interface, bci-processing-subsystem, sil-3, session-699, idempotency:ifc-bci-acq-to-artifact-699
IFC-BCIPROCESSINGSUBSYSTEM-003 | The interface between the Artifact Rejection Engine and the Feature Extraction Processor SHALL transfer cleaned 32-channel EEG epochs in 1-second windows with 50% overlap, formatted as float32 arrays, at a throughput of 2 epochs/second per channel. | interface-requirements | BCI Processing Subsystem |  | Test | 86 |  | 1-second window with 50% overlap is the standard for motor imagery classification: shorter windows sacrifice spectral resolution, longer windows add latency. Float32 format preserves dynamic range without overflow on standard DSP operations. 2 epochs/second output rate is the foundation for the 150ms pipeline latency in SUB-REQ-010. | interface, bci-processing-subsystem, sil-3, session-699, idempotency:ifc-bci-artifact-to-feat-699
IFC-BCIPROCESSINGSUBSYSTEM-004 | The interface between the Feature Extraction Processor and the BCI Classifier SHALL pass feature vectors containing CSP-projected band powers and SSVEP spectral amplitudes as float32 vectors, with timestamp and quality index, within 30ms of epoch receipt. | interface-requirements | BCI Processing Subsystem |  | Test | 86 |  | Feature vector format must include quality index (derived from electrode contact impedance) to allow the classifier to apply quality-weighted inference. 30ms is the allocation within the 150ms pipeline budget after 5ms acquisition and 20ms artifact rejection, leaving 50ms for classification and 45ms for arbitration. | interface, bci-processing-subsystem, sil-3, session-699, idempotency:ifc-bci-feat-to-classifier-699
IFC-BCIPROCESSINGSUBSYSTEM-005 | The interface between the BCI Classifier and the Command Arbitration Module SHALL pass command probability vectors for four navigation classes (forward, left, right, stop) as float32 with per-class confidence values, SNR index, and rolling 2-minute accuracy metric, at minimum 2Hz. | interface-requirements | BCI Processing Subsystem |  | Test | 86 |  | All four probability values required so the Command Arbitration Module can apply multi-class confidence thresholding and detect ambiguous dual-intent activations. Rolling accuracy metric enables SYS-REQ-008 adaptive mode transition without querying the classifier. SNR index enables SYS-REQ-007 signal-loss detection at the arbitration layer without feedback to the acquisition stage. | interface, bci-processing-subsystem, sil-3, session-699, idempotency:ifc-bci-classifier-to-arb-699
IFC-BCIPROCESSINGSUBSYSTEM-006 | The interface between the Command Arbitration Module and the Drive Subsystem SHALL use a CAN 2.0B message at 500 kbit/s, transmitting validated drive command (4 discrete values) and a 16-bit cyclic status counter, with maximum transmission latency of 5ms. | interface-requirements | BCI Processing Subsystem |  | Test | 71 |  | CAN 2.0B is selected for deterministic latency and EMI robustness in the motor-drive environment (12V bus noise). Status counter enables Drive Subsystem to detect missed arbitration frames and transition to safe stop. 5ms transmission latency is within the 150ms pipeline budget. CAN is consistent with the Drive Subsystem architecture decision (ARC-REQ-001). | interface, bci-processing-subsystem, sil-3, session-699, idempotency:ifc-bci-arb-to-drive-699
IFC-COMMUNICATIONSUBSYSTEM-001 | The EEG Accessibility Buggy SHALL communicate with the Companion App via BLE 5.0 LE Secure Connections (AES-CCM-128), providing status updates at ≥1 Hz and emergency notifications with ≤500ms latency. | interface-requirements | Communication Subsystem |  | Test | 100 |  | External companion app interface for carer situational awareness. BLE 5.0 LE Secure Connections (LESC) with AES-CCM-128 is the correct BLE 5.0 security mechanism (replaces the imprecise 'TLS 1.3 equivalent' phrasing). GDPR Article 9 requires link-layer encryption for EEG and health-status data. The ≥1 Hz status rate and ≤500ms emergency latency are derived from VER-REQ-109 acceptance criteria, which were validated against carer response-time requirements in the Signal Degradation scenario. | interface, external, communication, session-696, idempotency:ifc-ext-smartphone-696, red-team-session-715, reqs-eng-session-716
IFC-COMMUNICATIONSUBSYSTEM-002 | The interface between the EEG Accessibility Buggy and the Facility CMMS SHALL use a REST API over facility Wi-Fi (802.11ac) with TLS 1.3 encryption for automated maintenance log submission, fault reporting, and fleet telemetry, authenticated via API key or OAuth 2.0. | interface-requirements | Communication Subsystem |  | Test | 86 |  | External interface: CMMS integration automates compliance evidence for facility management. Wi-Fi provides sufficient bandwidth for periodic telemetry and maintenance logs. TLS 1.3 and authentication prevent unauthorised fleet data access and comply with GDPR for device usage data. | interface, external, communication, maintenance, session-696, idempotency:ifc-ext-cmms-696
IFC-COMMUNICATIONSUBSYSTEM-003 | The interface between the Communication Controller and the Bluetooth LE Module SHALL use UART at 1Mbps with hardware flow control (RTS/CTS), transferring HCI command and event packets per the Bluetooth HCI specification, with a maximum transfer latency of 2ms per HCI packet. | interface-requirements | Communication Subsystem |  | Test | 86 |  | HCI over UART at 1Mbps is the standard interface for Nordic nRF52840 host communication. Hardware flow control prevents buffer overflow at the 8kHz EEG data rate. 2ms maximum latency ensures the BLE connection interval of 7.5ms is not consumed by interface overhead. | interface, communication-subsystem, session-701, idempotency:ifc-comm-ctrl-to-ble-701
IFC-COMMUNICATIONSUBSYSTEM-004 | The interface between the Communication Controller and the Cellular Modem SHALL use USB 2.0 (CDC-ECM class) at 480Mbps with AT-command control channel, providing the Communication Controller with Ethernet-over-USB connectivity for LTE data at a maximum round-trip latency of 100ms to the LTE base station. | interface-requirements | Communication Subsystem |  | Test | 86 |  | The Communication Controller acts as data gateway and must command-and-control the Cellular Modem (APN config, registration, reset). USB CDC-ECM provides a standard IP stack interface avoiding proprietary serial AT tunnelling of IP frames. The 100ms RTT bound enables the 2-second telemetry uplink latency in SUB-REQ-046. | interface, communication-subsystem, session-703, idempotency:ifc-commctrl-cellmodem-703
IFC-COMMUNICATIONSUBSYSTEM-005 | The interface between the Bluetooth LE Module and the EEG headset SHALL use BLE 5.2 with a connection interval of 7.5ms, carrying 8-channel 250Hz EEG sample packets encoded in the EEG Headset Profile (proprietary GATT service) with end-to-end latency from electrode to BCI Acquisition Module not exceeding 20ms. | interface-requirements | Communication Subsystem |  | Test | 86 |  | The 50ms end-to-end BCI latency budget (SYS-REQ-001) allocates 20ms to wireless transport. BLE 5.2 at 7.5ms connection interval delivers 8-channel 16-bit EEG data (32 bytes per packet) within this window. Longer connection intervals cause acquisition dropouts; shorter intervals increase power consumption above the headset battery limit. | interface, communication-subsystem, session-703, idempotency:ifc-ble-headset-703
IFC-DRIVESUBSYSTEM-001 | The interface between the Motor Controller Unit and the Left Drive Motor Assembly SHALL consist of three-phase 20kHz PWM gate-drive signals at 48V and a 1024 PPR quadrature encoder feedback channel operating at a minimum sample rate of 10kHz, with an encoder signal propagation delay not exceeding 1ms. | interface-requirements | Drive Subsystem |  | Test | 86 |  | 20kHz PWM frequency is above the audible range to prevent tonal noise in clinical environments (IEC 60601-1 noise limits). 1024 PPR provides 0.77mm/pulse wheel displacement resolution, sufficient for the 0.1km/h velocity accuracy requirement. 10kHz encoder sampling (4× oversampling of 1024 PPR at max 2.5 rev/s) ensures Nyquist compliance. 1ms propagation limit fits within the 10ms velocity control cycle. | interface, drive-subsystem, session-700, sil-2, idempotency:ifc-mcu-leftmotor-700
IFC-DRIVESUBSYSTEM-002 | The interface between the Motor Controller Unit and the Right Drive Motor Assembly SHALL consist of three-phase 20kHz PWM gate-drive signals at 48V and a 1024 PPR quadrature encoder feedback channel operating at a minimum sample rate of 10kHz, with an encoder signal propagation delay not exceeding 1ms. | interface-requirements | Drive Subsystem |  | Test | 86 |  | Identical to the left motor interface specification — differential drive requires symmetric performance on both channels to achieve straight-line travel. Asymmetric encoder latency would cause directional drift under constant velocity commands. Matched interface specification simplifies integration test and EMC compliance (IEC 60601-1-2). | interface, drive-subsystem, session-700, sil-2, idempotency:ifc-mcu-rightmotor-700
IFC-DRIVESUBSYSTEM-003 | The interface between the Drive Power Stage and the Motor Controller Unit SHALL provide the 48V motor bus voltage to the gate driver circuits, return overcurrent trip status on a dedicated logic signal within 5ms of a trip event, and support a pre-charge sequencing signal from the Motor Controller Unit to enable soft-start of the 48V rail. | interface-requirements | Drive Subsystem |  | Test | 86 |  | The pre-charge sequencing prevents inrush current spikes when the Drive Subsystem powers on, protecting the LiFePO4 battery pack and motor MOSFET switches. Overcurrent trip feedback to the MCU in <5ms (matching the hardware trip time) ensures the MCU logs the event and can generate a diagnostic fault code before the watchdog triggers a safe-state. The 48V rail is isolated upstream by the Motor Power Isolation Relay, so the Power Stage–MCU interface is the last active interface before mechanical load. | interface, drive-subsystem, session-700, sil-2, idempotency:ifc-pwrstage-mcu-700
IFC-HMISUBSYSTEM-001 | The interface between the Main Application Processor and the Display Unit SHALL use SPI at 40MHz with a dedicated chip-select line, supporting a minimum frame rate of 30 fps at 800x480 resolution with 16-bit colour depth. | interface-requirements | HMI Subsystem |  | Test | 86 |  | 30 fps is sufficient for smooth status updates in a mobility aid context; higher frame rates are unnecessary and would increase EMI. SPI at 40MHz provides 320Mbps raw bandwidth, adequate for 800x480x16bit at 30fps (236Mbps required). | interface, hmi-subsystem, session-701, idempotency:ifc-map-display-701
IFC-HMISUBSYSTEM-002 | The interface between the Main Application Processor and the Audio Alert Module SHALL use I2C at 400kHz carrying alert-code commands with a command-to-onset latency not exceeding 50ms, where each alert code maps to a distinct tone pattern (frequency, duration, duty cycle) stored in the Audio Alert Module firmware. | interface-requirements | HMI Subsystem |  | Test | 86 |  | SYS-REQ-002 requires Emergency Stop actuation within 500ms of trigger detection. The 50ms audio alert onset budget is the human-perceptible notification latency allocated from the total response chain. I2C at 400kHz is sufficient for single-byte command codes; storing tone patterns in the Audio Alert Module firmware decouples HMI response from application processor availability during fault conditions. | interface, hmi-subsystem, session-703, idempotency:ifc-map-audio-703
IFC-HMISUBSYSTEM-003 | The interface between the Main Application Processor and the Status LED Array SHALL use SPI at 10MHz carrying RGB colour and brightness commands with a frame update latency not exceeding 20ms, supporting minimum 6 independently addressable LED elements. | interface-requirements | HMI Subsystem |  | Test | 86 |  | LED state must update within one control cycle (100ms) of a mode transition to avoid giving false state information to the care attendant. The 20ms SPI frame latency leaves 80ms for the application processor to detect the mode change and issue the update command. SPI is preferred over PWM for multi-element RGB addressing flexibility. | interface, hmi-subsystem, session-703, idempotency:ifc-map-ledarray-703
IFC-PERCEPTIONSUBSYSTEM-001 | The interface between the Forward Depth Sensor Array and the Perception MCU SHALL use I2C at 400 kHz (400 kHz I2C mode per UM10204) with three individually addressable sensor nodes, transferring an 8x8 depth grid per sensor at 10Hz with CRC-8 error detection on each transfer. | interface-requirements | Perception Subsystem |  | Test | 86 |  | I2C 400 kHz mode provides adequate bandwidth for 3x 64-byte depth grids at 10Hz (total 19.2 kbps, well within the 400 kHz bus capacity). Individual addressing allows the MCU to identify which sensor has failed. CRC-8 detects single-byte errors from EEG-band EMI as required by SYS-REQ-013. | interface, perception-subsystem, sil-2, session-701, idempotency:ifc-fwd-sensors-mcu-701
IFC-PERCEPTIONSUBSYSTEM-002 | The interface between the Side Proximity Sensor Pair and the Perception MCU SHALL use two dedicated GPIO lines (one per side), active-high TTL at 3.3V logic levels, with each sensor asserting the line within 30ms of detecting an object within 0.5m. | interface-requirements | Perception Subsystem |  | Test | 86 |  | GPIO is the lowest-latency interface option for a binary proximity alert and has no protocol overhead that could delay the alert beyond the 30ms window. Using dedicated lines per side allows the MCU to distinguish left and right obstacle sources for directional alert reporting. | interface, perception-subsystem, sil-2, session-701, idempotency:ifc-side-sensors-mcu-701
IFC-PERCEPTIONSUBSYSTEM-003 | The interface between the Perception MCU and the Safety Monitor Processor SHALL use SPI at 1MHz, transferring a 4-byte obstacle alert frame at 10Hz, containing a proximity status bitmask, minimum detected distance, and a rolling 8-bit frame counter, with the Safety Monitor Processor asserting a communication fault if two consecutive frames are missed. | interface-requirements | Perception Subsystem |  | Test | 86 |  | SPI is chosen over I2C for the safety interface because it provides deterministic transfer timing with no clock stretching ambiguity. A rolling frame counter allows the Safety Monitor Processor to detect missed frames independently of the communication line state, satisfying IEC 61508 SIL 2 independence requirements. The 4-byte frame is the minimum size to encode proximity status, distance, and counter. | interface, perception-subsystem, sil-2, session-701, idempotency:ifc-percmcu-smp-701
IFC-POWERSUBSYSTEM-001 | The interface between the EEG Accessibility Buggy and the Facility Charging Infrastructure SHALL use galvanically isolated IEC 60320 C13/C14 connectors or proprietary docking contacts, accepting 230V 50Hz or 120V 60Hz mains input, with drive functions locked during charging. | interface-requirements | Power Subsystem |  | Test | 86 |  | External interface: galvanic isolation protects the user from mains faults — the non-ambulatory user cannot self-disconnect. Drive lockout during charging prevents movement with a power cable connected. IEC 60601-1 requires double isolation for patient-connected medical devices. | interface, external, power, session-696, idempotency:ifc-ext-charging-696
IFC-POWERSUBSYSTEM-002 | The interface between the Battery Management System and the Safety Monitor Processor SHALL transmit a thermal fault signal within 50ms of any cell temperature exceeding 55°C, using a dedicated GPIO hardwired fault line with active-low logic, independent of CAN bus. | interface-requirements | Power Subsystem |  | Test | 71 |  | Dedicated GPIO fault line is required because the SIL-3 safety response to thermal runaway risk must not depend on CAN bus availability (which could be disrupted by the same electrical fault event). 55°C threshold is 5°C below the 60°C cutoff to give the Safety Monitor Processor a warning before hard disconnect, enabling graceful deceleration before power cutoff. | interface, power-subsystem, sil-3, session-699, idempotency:ifc-pwr-bms-to-safety-thermal-699
IFC-POWERSUBSYSTEM-003 | The interface between the Battery Management System and the Main Application Processor SHALL use CAN 2.0B at 250 kbit/s, transmitting pack voltage, individual cell voltages, SoC, temperature summary, and pack current at 100ms intervals. | interface-requirements | Power Subsystem |  | Test | 86 |  | 100ms update rate is adequate for SoC display refresh in HMI and charging state estimation. CAN 2.0B matches the vehicle bus standard. Cell-level voltage data required for compliance with IEC 62133 maintenance documentation and BMS diagnostic functions. SoC data drives the 4-hour runtime monitoring in SYS-REQ-004. | interface, power-subsystem, sil-3, session-699, idempotency:ifc-pwr-bms-to-main-can-699
IFC-SAFETYSUBSYSTEM-001 | The interface between the EEG Accessibility Buggy and the Facility Emergency/Nurse Call System SHALL transmit iBeacon-compatible BLE advertisements containing device ID and alert type within 500ms of Emergency Stop or seizure detection trigger, receivable by facility BLE gateways at a minimum range of 30m. | interface-requirements | Safety Subsystem |  | Test | 86 |  | External interface: the non-ambulatory user cannot call for help. The facility nurse call system is the fastest path to human assistance. 30m range covers typical hospital ward dimensions. iBeacon protocol ensures compatibility with existing facility BLE infrastructure. | interface, external, safety, communication, session-696, idempotency:ifc-ext-nursecall-696
IFC-SAFETYSUBSYSTEM-002 | The dedicated GPIO watchdog interface between the Main Application Processor (producer) and the Safety Monitor Processor (consumer) SHALL carry a 20ms-period square-wave heartbeat signal only, with no command or data payload on this line. | interface-requirements | Safety Subsystem |  | Test | 100 |  | IEC 61508-2 (Functional Safety of E/E/PE Safety-related Systems, Part 2) requires diagnostic test intervals ≤DC-category period for the claimed SIL. A 20ms heartbeat period (50Hz) provides 25 consecutive missed-pulse checks within the 500ms E-stop window specified in VER-REQ-004, improving fault-detection confidence over the previous 100ms period (5 missed pulses). The no-payload constraint preserves SMP independence — shared data channels between control and safety processors are prohibited under IEC 61508-2 clause 7.4.3. Dedicated GPIO eliminates shared-bus arbitration latency. | interface, safety-subsystem, session-697, idempotency:ifc-smp-map-heartbeat-697, red-team-session-715, reqs-eng-session-716
IFC-SAFETYSUBSYSTEM-003 | The interface between the Safety Monitor Processor and the Motor Power Isolation Relay SHALL use dual-channel opto-isolated control lines (one per relay coil), each driven by a dedicated GPIO output; both channels SHALL be required to be asserted simultaneously for relay energisation, such that single-channel failure results in relay de-energisation. | interface-requirements | Safety Subsystem |  | Test | 86 |  | Dual-channel opto-isolation prevents ground loops between the safety processor and the high-current relay circuit. The 1oo2 energisation architecture means either a stuck-at-low fault in one channel or a single opto-coupler failure causes safe-state relay opening — avoiding a single-point-of-failure that would allow the relay to remain closed on a fault. | interface, safety-subsystem, session-697, idempotency:ifc-smp-relay-control-697
IFC-SAFETYSUBSYSTEM-004 | The interface between the Inclinometer Tilt Sensor Unit and the Safety Monitor Processor SHALL be SPI at 1MHz, transferring a 6-byte data frame containing 16-bit pitch angle, 16-bit roll angle, and 16-bit CRC at 100Hz; the Safety Monitor Processor SHALL reject frames with CRC errors and SHALL trigger Emergency Stop if 5 consecutive frames are rejected. | interface-requirements | Safety Subsystem |  | Test | 86 |  | SPI is selected over I2C for deterministic transfer latency — I2C clock stretching can violate the 100Hz data rate budget. CRC and consecutive-error E-stop protect against sensor communication faults being silently treated as valid zero-tilt readings, which would mask an actual tip-over condition. | interface, safety-subsystem, session-697, idempotency:ifc-incl-smp-spi-697
IFC-SAFETYSUBSYSTEM-005 | The interface between the Seizure Detection Module and the Safety Monitor Processor SHALL be a shared memory region on the Safety Monitor Processor, updated by the Seizure Detection Module at 50Hz with a 1-byte status word (bit 0 = seizure flag, bit 1 = algorithm health), protected by a mutex; the Safety Monitor Processor SHALL read this region within 10ms of each update. | interface-requirements | Safety Subsystem |  | Test | 86 |  | Seizure Detection Module executes as a co-resident software task on the Safety Monitor Processor — inter-task shared memory with mutex is lower latency than inter-processor communication (which would require a serial bus with handshake overhead). 50Hz update rate provides 20ms temporal resolution — adequate since the detection algorithm requires 150ms pattern window, meaning the Safety Monitor Processor will see the flag within one read cycle of it being set. | interface, safety-subsystem, session-697, idempotency:ifc-sdm-smp-shmem-697
IFC-VEHICLEPLATFORM-001 | The interface between the Electronics Bay and the Chassis Frame SHALL use four M6 stainless steel fasteners with self-locking nuts, providing a pull-out strength of at least 500N per fastener, and shall include a rubber vibration isolator rated for 3g shock attenuation to protect electronic components during kerb crossing. | interface-requirements | Vehicle Platform |  | Test | 71 |  | 500N pull-out strength exceeds the 3g shock force applied to the 2kg Electronics Bay mass by a factor of 8, providing adequate margin for ISO 7176-8 compliance. Rubber isolation prevents resonant transmission of kerb-crossing vibration into the processor PCBs which are rated for 2g sinusoidal vibration. | interface, vehicle-platform, session-701, idempotency:ifc-ebay-chassis-701
IFC-VEHICLEPLATFORM-002 | The interface between the Wheel and Caster Assembly and the Chassis Frame SHALL use a tool-free quick-release mechanism requiring a maximum 15N force for removal and replacement, with a mechanical interlock preventing partial engagement (either fully locked or fully detached). | interface-requirements | Vehicle Platform |  | Test | 86 |  | Wheel replacement is a field maintenance task performed by care facility staff without specialist tools (SYS-REQ-015). The 15N force limit is derived from ISO 7176-1 control force limits adapted for maintenance operations — ensuring the task is within capability of care staff regardless of hand strength. The binary interlock prevents wheel detachment due to partial engagement during operation, which would be a safety-critical failure. | interface, vehicle-platform, session-703, idempotency:ifc-wheel-chassis-703
STK-REQ-001 | The EEG Accessibility Buggy SHALL enable the user to navigate independently between locations within a care facility and along outdoor paved paths using only EEG brain signals, without requiring continuous attendant assistance. | stakeholder-requirements | Stakeholder Requirements |  | Demonstration | 100 |  | BCI User, Daily Independent Navigation scenario: the primary purpose of the system is to restore independent mobility to individuals with severe motor disabilities who cannot operate conventional wheelchairs or joystick-controlled devices. | stakeholder, stk-bci-user, session-696, idempotency:stk-user-independent-nav-696
STK-REQ-002 | The EEG Accessibility Buggy SHALL protect the user from injury during all operating modes, including autonomous collision avoidance, emergency stopping within 200ms of safety trigger detection, and automatic seizure response. | stakeholder-requirements | Stakeholder Requirements |  | Test | 100 |  | BCI User, all scenarios: user has severe motor disability and cannot self-extract from hazardous situations — the system is the sole protection layer between cognitive intent and physical harm. Driven by H-001 (signal loss), H-002 (collision), H-006 (seizure). | stakeholder, stk-bci-user, safety, session-696, idempotency:stk-user-safety-696
STK-REQ-003 | The EEG Accessibility Buggy SHALL detect declining cognitive performance and adapt its operating mode to prevent dangerous loss of BCI control accuracy, alerting the user and carer before accuracy drops below safe thresholds. | stakeholder-requirements | Stakeholder Requirements |  | Test | 100 |  | BCI User, Cognitive Fatigue scenario: after 25-30 minutes of continuous BCI use, classification accuracy degrades from 85% to 68%. Without detection and adaptation, misclassified commands become a collision and safety risk (H-004). | stakeholder, stk-bci-user, session-696, idempotency:stk-user-fatigue-696
STK-REQ-004 | The EEG Accessibility Buggy SHALL provide a care attendant manual override operable by untrained personnel within 60 seconds of a ≤5-minute orientation, transferring full steering and throttle authority within 100ms while maintaining obstacle detection. | stakeholder-requirements | Stakeholder Requirements |  | Demonstration | 100 |  | Care Attendant scenario: carer rotation means individuals may encounter the device with no prior training. The ≤5-minute orientation threshold reflects NHS Band 3 carer induction constraints (maximum 5-minute device briefing per Assistive Technology Support Protocol). The 60-second operational readiness criterion is derived from emergency response time requirements — if a clinical incident occurs, the carer must be able to take control immediately after a brief orientation. 100ms handover latency ensures no perceptible delay during the authority transfer. Obstacle detection remains active per BS EN ISO 13482 (Safety requirements for personal care robots) to prevent collision hazards during carer operation. | stakeholder, stk-carer, session-696, idempotency:stk-carer-override-696, red-team-session-715, reqs-eng-session-716
STK-REQ-005 | The EEG Accessibility Buggy SHALL generate a ≥75 dBA audible alert and a colour-coded LED distinguishable at ≥5 m (amber: Degraded mode, red: Emergency Stop) within 1 second of any mode-transition event. | stakeholder-requirements | Stakeholder Requirements |  | Test | 100 |  | Care Attendant scenario: carer may be in an adjacent room or attending another patient. The ≥75 dBA threshold at 1 m (measured per IEC 60068-2-64: Vibration test, acoustic emission) is the minimum level audible over typical clinical environment noise (NHS estates guidance: 65 dB background). The ≥5 m LED distinguishability criterion is derived from typical ward bay dimensions (4 m × 6 m) — a carer at the far end of the bay must see the alert. Amber/red colour coding follows BS EN 60073 (Basic safety principles for man-machine interfaces) actuator/indicator assignment. The 1-second response deadline ensures the alert precedes any secondary clinical consequence of the mode transition. | stakeholder, stk-carer, session-696, idempotency:stk-carer-alerts-696, red-team-session-715, reqs-eng-session-716
STK-REQ-006 | The EEG Accessibility Buggy SHALL provide a wired diagnostic interface enabling the maintenance technician to execute a complete automated safety-critical subsystem test suite within 20 minutes per buggy. | stakeholder-requirements | Stakeholder Requirements |  | Demonstration | 100 |  | Maintenance Technician, Weekly Maintenance scenario: fleet of 4 buggies must be inspected in a 2-hour maintenance window. 20 min per buggy (80 min total) leaves margin for fault investigation and electrode replacement. | stakeholder, stk-maint-tech, session-696, idempotency:stk-maint-diagnostics-696
STK-REQ-007 | The EEG Accessibility Buggy SHALL support over-the-air and wired firmware updates with automatic rollback to the last-verified firmware version if the update fails post-flash verification, installable by a trained facility technician. | stakeholder-requirements | Stakeholder Requirements |  | Test | 100 |  | Maintenance Technician, Weekly Maintenance scenario: firmware updates are applied weekly from the facility server. Dual-bank flash with rollback is essential because a bricked buggy removes it from the fleet and impacts user mobility. IEC 62304 requires controlled software update processes. | stakeholder, stk-maint-tech, session-696, idempotency:stk-maint-firmware-696
STK-REQ-008 | The EEG Accessibility Buggy SHALL allow the clinical prescriber to configure patient-specific BCI parameters including speed limits, available command palette, maximum session duration, and classification sensitivity thresholds via a secure authenticated interface. | stakeholder-requirements | Stakeholder Requirements |  | Demonstration | 100 |  | Clinical Prescriber (OT/rehab physician): each user has different motor imagery capability, fatigue tolerance, and risk profile. A C4 SCI patient may need different parameters than an ALS patient. Clinical configuration drives safe individualised therapy. | stakeholder, stk-clinical, session-696, idempotency:stk-clinical-config-696
STK-REQ-009 | The EEG Accessibility Buggy SHALL record session data (duration, BCI accuracy, fatigue events, mode transitions, distance) and export it as JSON via the maintenance interface within 60 seconds of session end. | stakeholder-requirements | Stakeholder Requirements |  | Test | 100 |  | Clinical Prescriber requirement: outcome data must be exportable for NHS clinical audit trails and therapy effectiveness review. JSON format is specified because it is machine-readable by standard clinical outcome tools and is required by NHS Digital Data Standards for Class IIb medical device reporting. The maintenance interface (USB-C service port, per IFC-REQ-024) is the designated export channel to prevent unauthorised BLE data extraction. The 60-second export window ensures data availability before the next session begins. | stakeholder, stk-clinical, session-696, idempotency:stk-clinical-data-696, red-team-session-715, reqs-eng-session-716
STK-REQ-010 | The EEG Accessibility Buggy SHALL provide a complete design history file, risk management file, and clinical evaluation report compliant with IEC 60601-1 (Medical electrical equipment — General requirements for basic safety and essential performance), IEC 62304 (Medical device software — Software life cycle processes), ISO 14971 (Application of risk management to medical devices), and EN 12184 (Electrically powered wheelchairs, scooters and their chargers — Requirements and test methods). | stakeholder-requirements | Stakeholder Requirements |  | Inspection | 86 |  | Regulatory Authority (MHRA/FDA/TGA): the system is a Class IIb medical device. Without compliant documentation, the device cannot receive market authorisation and cannot legally be used with patients. | stakeholder, stk-regulatory, session-696, idempotency:stk-reg-compliance-696
STK-REQ-011 | The EEG Accessibility Buggy SHALL support post-market surveillance data collection including automated incident logging, usage statistics aggregation, adverse event reporting, and software version tracking across the deployed fleet. | stakeholder-requirements | Stakeholder Requirements |  | Inspection | 100 |  | Regulatory Authority: EU MDR Article 83 and FDA 21 CFR 803 require active post-market surveillance for Class IIb devices. Automated data collection reduces reporting burden and ensures adverse events are captured promptly. | stakeholder, stk-regulatory, session-696, idempotency:stk-reg-postmarket-696
STK-REQ-012 | The EEG Accessibility Buggy SHALL integrate with the facility's Computerised Maintenance Management System (CMMS) via REST API for automated maintenance scheduling, fault reporting, and fleet utilisation tracking. | stakeholder-requirements | Stakeholder Requirements |  | Test | 100 |  | Facility Management: facilities operate fleets of 2-8 buggies. Manual maintenance tracking is error-prone — missed inspections create regulatory and safety liability. CMMS integration automates compliance evidence. | stakeholder, stk-facility, session-696, idempotency:stk-facility-fleet-696
STK-REQ-013 | The EEG Accessibility Buggy SHALL operate within existing facility infrastructure without requiring structural modifications beyond standard accessibility provisions compliant with the Disability Discrimination Act and Building Regulations. | stakeholder-requirements | Stakeholder Requirements |  | Analysis | 100 |  | Facility Management: care facilities and hospitals cannot undertake corridor widening or door replacement for a mobility device. The buggy must fit existing DDA-compliant corridors (≥1200mm), doorways (≥900mm), and lifts (≥1100×1400mm). | stakeholder, stk-facility, session-696, idempotency:stk-facility-infra-696
STK-REQ-014 | The EEG Accessibility Buggy SHALL maintain safe stopping distances and exhibit predictable, smooth movement patterns to avoid startling or injuring bystanders sharing indoor corridors and outdoor pathways. | stakeholder-requirements | Stakeholder Requirements |  | Test | 100 |  | Bystanders/Pedestrians, collision scenario (H-002): bystanders include elderly residents with slow reaction times, children, and wheelchair users. Unpredictable movement or abrupt stops create secondary injury risk in shared spaces. | stakeholder, stk-bystander, session-696, idempotency:stk-bystander-safety-696
STK-REQ-015 | The EEG Accessibility Buggy SHALL provide audible low-speed travel tones and visible direction indicators to alert nearby pedestrians of the vehicle's presence and intended direction of travel in shared indoor and outdoor spaces. | stakeholder-requirements | Stakeholder Requirements |  | Test | 100 |  | Bystanders/Pedestrians: unlike conventional wheelchairs, a BCI-controlled buggy gives no visible human-body cues about direction. Pedestrians cannot predict movement from the user's gaze or posture, requiring engineered warning signals. | stakeholder, stk-bystander, session-696, idempotency:stk-bystander-warning-696
STK-REQ-016 | The EEG Accessibility Buggy SHALL maintain safe operation in electromagnetically noisy environments including hospital corridors with MRI fringe fields (5 gauss line), electrosurgical units, diathermy equipment, and 50Hz fluorescent lighting ballast harmonics. | stakeholder-requirements | Stakeholder Requirements |  | Test | 100 |  | Environmental constraint, Signal Degradation scenario and H-007: hospital and care facility environments contain significant EMI sources that directly interfere with EEG signal acquisition. The system must degrade safely rather than misclassify corrupted signals. | stakeholder, stk-environment, session-696, idempotency:stk-env-emi-696
STK-REQ-017 | The EEG Accessibility Buggy SHALL operate safely on outdoor paved paths in ambient temperatures from -5°C to 40°C, with rain exposure protection to IP44 minimum for all outdoor-rated components, and on gradients up to 8 degrees. | stakeholder-requirements | Stakeholder Requirements |  | Test | 86 |  | Environmental constraint, Daily Independent Navigation scenario: users navigate between indoor and outdoor spaces (e.g., care home to garden patio). The system must handle UK/temperate climate conditions without degradation of safety functions. | stakeholder, stk-environment, session-696, idempotency:stk-env-outdoor-696
STK-REQ-018 | The EEG Accessibility Buggy SHALL handle, store, and transmit EEG biometric data in compliance with GDPR and applicable data protection regulations, including encryption at rest and in transit, user consent management, and data retention policies. | stakeholder-requirements | Stakeholder Requirements |  | Inspection | 100 |  | Regulatory Authority and BCI User: EEG data is biometric personal data under GDPR Article 9. Brain signal recordings reveal health status and potentially cognitive state — requiring the highest category of data protection. | stakeholder, stk-regulatory, stk-bci-user, session-696, idempotency:stk-data-protection-696
SUB-BCIPROCESSINGSUBSYSTEM-001 | The Artifact Rejection Engine SHALL reject electromagnetic and muscle artefact from 32-channel EEG within 20ms of epoch receipt, with false rejection of true neural signals not exceeding 10%. | subsystem-requirements | BCI Processing Subsystem |  | Test | 100 |  | 20ms latency budget derived from 50ms total BCI pipeline budget (SYS-REQ-001 <150ms classification latency). False rejection of >10% degrades classification accuracy below the 70% threshold in SYS-REQ-008, causing unnecessary safe-stop transitions. IEC 60601-1 applies to EMC performance of medical devices. | subsystem, bci-processing-subsystem, sil-3, session-699, idempotency:sub-bci-artifact-rejection-699
SUB-BCIPROCESSINGSUBSYSTEM-002 | The BCI Classifier SHALL achieve a minimum four-class motor imagery and SSVEP classification accuracy of 75% across a 2-minute rolling window during normal operation. | subsystem-requirements | BCI Processing Subsystem |  | Test | 100 |  | 75% accuracy floor is derived from SYS-REQ-008 which triggers adaptive feedback mode at 70%. The 5% margin allows the Command Arbitration Module's confidence threshold (70% per-command) to remain effective without generating spurious safe-stop events under mildly degraded signal conditions. Below 70% system accuracy the user loses meaningful volitional control, creating an unsafe condition. | subsystem, bci-processing-subsystem, sil-3, session-699, idempotency:sub-bci-classifier-accuracy-699
SUB-BCIPROCESSINGSUBSYSTEM-003 | The Command Arbitration Module SHALL emit a validated drive command to the Drive Subsystem interface within 150ms of the corresponding EEG signal epoch being acquired by the EEG Acquisition Module. | subsystem-requirements | BCI Processing Subsystem |  | Test | 100 |  | 150ms end-to-end latency directly derived from SYS-REQ-001. This is the cognitive-motor feedback loop constraint: delays >200ms cause users to generate corrective commands before the original command executes, resulting in command accumulation and erratic vehicle behaviour. 150ms allows 50ms margin below the 200ms perceptual threshold. | subsystem, bci-processing-subsystem, sil-3, session-699, idempotency:sub-bci-arbitration-latency-699
SUB-BCIPROCESSINGSUBSYSTEM-004 | When BCI signal-to-noise ratio remains below the classification threshold for more than 3 consecutive seconds, the Command Arbitration Module SHALL emit a STOP command to the Drive Subsystem and suppress further drive commands until SNR recovers. | subsystem-requirements | BCI Processing Subsystem |  | Test | 86 |  | Directly implements SYS-REQ-007 at the subsystem level. 3-second timeout balances false-positive rate (transient signal dropout from head movement) against safety risk of continued motion with unreliable control. Safe state is vehicle stop, consistent with the hazard register. SIL-3 requirement because failure to stop on signal loss could result in uncontrolled vehicle motion near vulnerable users. | subsystem, bci-processing-subsystem, sil-3, session-699, idempotency:sub-bci-signal-loss-safe-state-699
SUB-BCIPROCESSINGSUBSYSTEM-005 | The Feature Extraction Processor SHALL load and apply per-user CSP spatial filter matrices from encrypted calibration storage within 5 seconds of user session initialisation. | subsystem-requirements | BCI Processing Subsystem |  | Test | 100 |  | MoP basis: Published motor imagery BCI research (Blankertz et al. 2008, 'Optimizing Spatial Filters for Robust EEG Single-Trial Analysis', NeuroImage) documents 15–25% accuracy reduction when using population-averaged vs. subject-specific CSP spatial filters across EEG motor imagery tasks. Per-user matrices are therefore essential for the ≥80% classification accuracy required by SYS-REQ-006. 5-second load time is within the SYS-REQ-018 30-second startup budget. Encryption required under STK-REQ-018 (EEG biometric data, GDPR Article 9 special category). Failure to load forces operation on default matrices, triggering SYS-REQ-007 signal degradation pathway and Degraded/Assisted mode. | subsystem, bci-processing-subsystem, sil-3, session-699, idempotency:sub-bci-calibration-load-699, reqs-eng-session-716
SUB-BCIPROCESSINGSUBSYSTEM-006 | The Artifact Rejection Engine SHALL expose a watchdog interface to the BCI Processing Subsystem supervisor; when no valid epoch output is received for more than 500ms, the supervisor SHALL reset the Artifact Rejection Engine and log the event for post-session analysis. | subsystem-requirements | BCI Processing Subsystem |  | Test | 86 |  | IEC 61508 SIL-3 requires monitoring of autonomous processing functions. The Artifact Rejection Engine is functionally autonomous (runs its own ICA pipeline without per-frame supervision) but can enter a deadlock or memory leak state. 500ms without output is an unambiguous failure indicator — the nominal output period is 500ms (2 epochs/second), so one missed window triggers reset. This provides a bounded recovery time compatible with the BCI pipeline latency budget. | subsystem, bci-processing-subsystem, sil-3, session-699, idempotency:sub-bci-artifact-engine-watchdog-699
SUB-BCIPROCESSINGSUBSYSTEM-007 | The Main Application Processor SHALL expose a USB-C service port accessible on the Electronics Bay panel that provides a diagnostic test suite interface, executing the full motor response, brake torque, sensor calibration, battery cell health, and BCI pipeline integrity test sequence within 20 minutes via authenticated USB connection. | subsystem-requirements | BCI Processing Subsystem |  | Test | 86 |  | SYS-REQ-015 requires a USB-C diagnostic suite completing in 20 minutes. The Main Application Processor hosts the diagnostic executive and has access to all subsystem buses (CAN, I2C, SPI, UART). Authentication is required to prevent unauthenticated firmware or calibration access. 20-minute completion time is derived from scheduled maintenance window constraints in care facility operational planning. | subsystem, bci-processing-subsystem, session-702, idempotency:sub-map-usb-service-port-702
SUB-BCIPROCESSINGSUBSYSTEM-008 | The Main Application Processor SHALL encrypt all EEG biometric data using AES-256 at rest and TLS 1.3 in transit, enforce role-based access control with at least three roles (clinical, maintenance, user), and maintain an append-only audit log of all data access events with timestamp, user identity, and operation type. | subsystem-requirements | BCI Processing Subsystem |  | Test | 86 |  | SYS-REQ-019 mandates AES-256 and TLS 1.3 for EEG biometric data. The Main Application Processor is the data custodian for all BCI pipeline data. Three-role RBAC (clinical, maintenance, user) is the minimum partition required by GDPR Article 25 (data protection by design) and IEC 62443 (Security for industrial automation) tier partitioning. Audit logging supports forensic investigation and regulatory compliance. | subsystem, bci-processing-subsystem, security, session-702, idempotency:sub-map-aes256-security-702
SUB-BCIPROCESSINGSUBSYSTEM-009 | The Artifact Rejection Engine SHALL operate from the 3.3V logic rail supplied by the DC-DC Converter Array, with a maximum peak current draw of 250mA during active signal processing, and SHALL maintain operation without error for supply voltage variation between 3.0V and 3.6V. | subsystem-requirements | BCI Processing Subsystem |  | Test | 86 |  | Lint finding: artifact rejection engine is Powered but has no power source requirements. The ARE is a software-firmware block executing on the BCI DSP and draws power through the DSP supply rail. Bounding the current draw to 250mA peak prevents thermal shutdown of the 3.3V converter during simultaneous EEG acquisition and artifact processing bursts. Brown-out on this rail would cause BCI pipeline failure and trigger emergency stop. | subsystem, bci-processing-subsystem, power, sil-3, session-702, idempotency:sub-bci-artifact-engine-power-702
SUB-BCIPROCESSINGSUBSYSTEM-010 | While in Normal Navigation mode, the BCI Processing Subsystem SHALL provide a software watchdog that monitors BCI command output validity and SHALL halt command generation and assert a safe-stop signal within 100ms if command outputs deviate from expected bounds or if the watchdog is not kicked within a 200ms window. | subsystem-requirements | BCI Processing Subsystem |  | Test | 86 |  | Lint finding: Normal Navigation mode is a Functionally Autonomous operational state with no safety or override constraints. This requirement addresses the autonomous operation safety gap per IEC 61508 (Functional safety of E/E/PE safety-related systems) SIL-3 requirements for autonomous command generation. The 100ms halt time is bounded by the 200ms total emergency stop response budget from SYS-REQ-002. | subsystem, bci-processing-subsystem, safety, sil-3, session-702, idempotency:sub-bci-normal-nav-watchdog-702
SUB-BCIPROCESSINGSUBSYSTEM-011 | The EEG Accessibility Buggy electrode interface components that contact the user scalp SHALL be manufactured from materials meeting ISO 10993-1 (Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process) biocompatibility requirements for skin-contacting medical devices, and SHALL support decontamination with standard clinical disinfectants between users without degradation of electrical performance. | subsystem-requirements | BCI Processing Subsystem |  | Analysis | 86 |  | The buggy is Biological/Biomimetic (interfaces with biological EEG signals via wet electrodes on user scalp). Scalp-contact electrodes are skin-contacting medical devices under ISO 10993-1. Clinical deployment in care facilities serving multiple users requires decontamination capability to prevent infection transmission. IEC 60601-1 (Medical electrical equipment — Part 1: General requirements for basic safety) also requires biocompatibility for patient-applied parts. | subsystem, bci-processing-subsystem, regulatory, biocompatibility, session-702, idempotency:sub-bci-electrode-biocompatibility-702
SUB-BCIPROCESSINGSUBSYSTEM-012 | The Artifact Rejection Engine SHALL execute as a software module on the Feature Extraction Processor hardware, occupying no more than 30% of available CPU cycles during continuous EEG processing at 250 SPS per channel, with a deterministic maximum latency of 4ms per processing frame. | subsystem-requirements | BCI Processing Subsystem |  | Test | 86 |  | The lint finding flags the Artifact Rejection Engine as lacking Physical Object trait but having physical embodiment constraints. As a software algorithm running on the Feature Extraction Processor, it must be bounded in resource consumption to not starve classifier and acquisition processes. The 30% CPU budget preserves headroom for the BCI Classifier (50%) and overhead (20%). The 4ms latency ensures artifact-free EEG is available within the 10ms BCI processing frame. | subsystem, bci-processing-subsystem, sil-3, session-707, idempotency:sub-bci-artifact-hw-context-707
SUB-BCIPROCESSINGSUBSYSTEM-013 | The Main Application Processor SHALL validate all BCI-derived navigation commands using a cryptographic message authentication code (HMAC-SHA256) when received from any external software interface, and SHALL reject and log any command that fails authentication, missing a valid session token, or arrives via an interface not in the authorised interface table defined during commissioning. | subsystem-requirements | BCI Processing Subsystem |  | Test | 86 |  | Normal Navigation mode and Artifact Rejection Engine are classified as Digital/Virtual with no cybersecurity requirement at SUB level. A BCI mobility device controls physical motion — unauthorised injection of navigation commands could result in collision or unsafe motion. HMAC-SHA256 command authentication provides integrity protection against replay attacks and command injection without requiring a symmetric key exchange per session, which is appropriate for the embedded MAP execution environment. This requirement closes the cybersecurity gap identified by UHT Substrate classification of the Normal Navigation mode (hex 40B72300). | subsystem, bci-processing, security, session-708, qc, idempotency:sub-map-command-authentication-708
SUB-COMMUNICATIONSUBSYSTEM-001 | The Bluetooth LE Module SHALL maintain a stable BLE 5.2 connection to the paired EEG headset at 2.4GHz with a connection interval of 7.5ms, sustaining 32-channel EEG data at 8kHz sample rate (effective throughput 512kbps) within an operating range of 3m with no physical obstructions. | subsystem-requirements | Communication Subsystem |  | Test | 86 |  | BCI signal path requires continuous EEG data at 8kHz to meet the classification accuracy in SYS-REQ-001. The 7.5ms BLE connection interval is the minimum supported by BLE 5.2 and the tightest interval that avoids BCI latency degradation. 3m range covers all typical headset-to-buggy cable routing scenarios within a clinical environment. | subsystem, communication-subsystem, sil-1, session-701, idempotency:sub-comm-ble-throughput-701
SUB-COMMUNICATIONSUBSYSTEM-002 | The Communication Controller SHALL enforce a firewall rule that prevents all data traffic originating from the Cellular Modem from reaching the internal CAN bus or I2C/SPI control networks. | subsystem-requirements | Communication Subsystem |  | Inspection | 86 |  | A cellular modem is an externally accessible network endpoint. Without isolation, a compromised modem could inject spoofed drive or safety commands on the CAN bus, constituting a cybersecurity attack path directly threatening occupant safety. The firewall is the primary control for IEC 62443 network segmentation of the safety domain from the external domain. | subsystem, communication-subsystem, sil-1, session-701, idempotency:sub-comm-firewall-isolation-701
SUB-COMMUNICATIONSUBSYSTEM-003 | When Emergency Stop or seizure detection triggers, the Communication Controller SHALL transmit a BLE 5.0 iBeacon-compatible alert to the facility BLE gateway within 500ms, including device ID, alert type, and GPS coordinates, and simultaneously notify the paired caregiver smartphone. | subsystem-requirements | Communication Subsystem |  | Test | 86 |  | SYS-REQ-017 requires BLE alert transmission within 500ms on emergency triggers. The Communication Controller manages the BLE 5.0 module (nRF52840) and is the component responsible for alert generation. 500ms is the maximum acceptable caregiver notification latency before a seizure or emergency becomes unmanageable. | subsystem, communication-subsystem, sil-1, session-702, idempotency:sub-comm-ble-alert-702
SUB-COMMUNICATIONSUBSYSTEM-004 | The Cellular Modem SHALL authenticate to the remote telemetry server using mutual TLS 1.3 with a device-unique X.509 certificate provisioned at manufacture, rejecting any server certificate not signed by the system's root CA. | subsystem-requirements | Communication Subsystem |  | Test | 100 |  | SYS-REQ-019 requires TLS 1.3 for data in transit. The Cellular Modem is the only component with external IP connectivity; mutual TLS with per-device certificates prevents impersonation of the server (downgrade attack) and spoofing of the device. Device certificates provisioned at manufacture ensure the credential chain cannot be forged in the field. | subsystem, communication-subsystem, session-703, idempotency:sub-comm-mtls-703
SUB-COMMUNICATIONSUBSYSTEM-005 | The Cellular Modem SHALL transmit session telemetry (location, BCI state, battery, error events) to the remote server at a minimum interval of 10 seconds with a maximum uplink latency of 2 seconds under nominal 4G coverage. | subsystem-requirements | Communication Subsystem |  | Test | 86 |  | Care attendants and clinical teams rely on near-real-time session data for remote supervision of users with severe motor impairments. A 10-second interval balances cellular data consumption against clinical oversight needs. The 2-second uplink latency bound ensures that location and safety events reach the monitoring platform while the user is still at the reported position. | subsystem, communication-subsystem, session-703, idempotency:sub-comm-telemetry-rate-703
SUB-COMMUNICATIONSUBSYSTEM-006 | The Communication Subsystem SHALL provide a USB-C service port that exposes a diagnostic API enabling technicians to execute motor response verification tests and receive pass/fail results with logged timestamps within 10 s of command initiation. | subsystem-requirements | Communication Subsystem |  | Test | 100 |  | SYS-REQ-015 specifies USB-C access to an automated diagnostic suite; the Communication Subsystem hosts the USB-C physical interface and routes diagnostic commands to the Main Application Processor, making it the correct implementation locus. | subsystem, communication-subsystem, session-705, idempotency:sub-comm-usbc-diag-705
SUB-COMMUNICATIONSUBSYSTEM-007 | When a Seizure Emergency Stop event is triggered, the Communication Controller SHALL transmit a BLE 5.0 emergency alert packet containing device ID, alert type code, GPS coordinates (WGS-84 decimal), and UTC timestamp to the facility emergency system receiver within 500ms of the Emergency Stop command. | subsystem-requirements | Communication Subsystem |  | Test | 86 |  | Derives from SYS-REQ-005 which requires BLE emergency alert transmission to facility emergency system on seizure detection. 500ms transmission budget allows for BLE connection establishment (max 200ms for pre-paired devices) plus packet delivery, while the Emergency Stop itself is completed within 150ms. The facility receiver must get actionable location data to dispatch care staff rapidly. | subsystem, communication-subsystem, sil-1, session-707, idempotency:sub-comm-ble-emergency-alert-707
SUB-DRIVESUBSYSTEM-001 | The Motor Controller Unit SHALL implement closed-loop velocity control for each drive motor channel, updating the velocity setpoint at 100 Hz using 1024 PPR quadrature encoder feedback, with a steady-state velocity error not exceeding ±0.1 km/h at any commanded speed. | subsystem-requirements | Drive Subsystem |  | Test | 86 |  | Closed-loop control at 100Hz is required to achieve accurate differential-drive steering. The ±0.1 km/h accuracy threshold is derived from SYS-REQ-003 speed limits: at 6km/h maximum, a ±1.7% error is acceptable for stable path tracking while avoiding overspeed. Encoder feedback rate of 1024 PPR at 250mm wheel diameter gives 0.77mm/pulse resolution, sufficient for the control loop. | subsystem, drive-subsystem, session-700, sil-2, idempotency:sub-mcu-velocity-control-700
SUB-DRIVESUBSYSTEM-002 | The Motor Controller Unit SHALL enforce a vehicle speed limit of 6 km/h in Normal Navigation mode and 2 km/h in Restricted mode, clamping any velocity command exceeding these thresholds before the command is applied to the motor drive stages. | subsystem-requirements | Drive Subsystem |  | Test | 86 |  | Derives from SYS-REQ-003. Speed limits are enforced in firmware to prevent unsafe motion: 6km/h is consistent with IEC 62133 electric mobility device guidance for indoor environments, and 2km/h Restricted mode is for narrow corridors and near-bystander situations. Clamping in MCU firmware rather than relying on upstream command rejection provides defence-in-depth at SIL 2. | subsystem, drive-subsystem, session-700, sil-2, idempotency:sub-mcu-speed-limit-700
SUB-DRIVESUBSYSTEM-003 | When the Motor Controller Unit receives a kill signal via the CAN safety frame or detects loss of CAN heartbeat for more than 100ms, the Motor Controller Unit SHALL apply regenerative braking to bring both motors to zero velocity within 250ms and hold zero-velocity command until a valid restart sequence is received. | subsystem-requirements | Drive Subsystem |  | Test | 71 |  | Derives from SYS-REQ-002: 250ms safe stop is the system-level budget. The MCU's 250ms hard stop matches the system-level requirement. CAN heartbeat loss at 100ms provides early detection (50% of the 200ms upstream timeout in SUB-REQ-010) to allow regenerative braking time. Regenerative braking is preferred over freewheeling to maintain directional control during deceleration on slopes. | subsystem, drive-subsystem, session-700, sil-2, safety, idempotency:sub-mcu-safe-state-700
SUB-DRIVESUBSYSTEM-004 | The Drive Power Stage SHALL trip the hardware overcurrent protection circuit and isolate the motor phase outputs within 5ms when per-channel motor current exceeds 30A, independent of Motor Controller Unit firmware. | subsystem-requirements | Drive Subsystem |  | Test | 100 |  | Hardware overcurrent protection independent of MCU firmware is required at SIL 2 because firmware-only protection cannot be credited under IEC 61508 (Functional safety of E/E/PE safety-related systems) at SIL 2 without redundant validation. 30A trip threshold is derived from 500W peak motor rating at 48V nominal (10.4A max continuous, 30A provides 3x margin for inrush while protecting MOSFET switches rated at 80V/40A). 5ms trip time ensures MOSFET junctions do not exceed thermal limits. | subsystem, drive-subsystem, session-700, sil-2, safety, idempotency:sub-drive-pwr-overcurrent-700
SUB-DRIVESUBSYSTEM-005 | The Left Drive Motor Assembly and the Right Drive Motor Assembly SHALL each provide a minimum continuous shaft output of 250W at 48V DC supply voltage across the operating temperature range of 0°C to 40°C, with the driven wheel maintaining traction on level indoor surfaces up to a 5% gradient. | subsystem-requirements | Drive Subsystem |  | Test | 86 |  | 250W continuous per motor is required to achieve 6km/h with a 150kg total vehicle load (occupant + buggy) on a 5% gradient, per the mechanical traction analysis in the ConOps. 48V supply matches the LiFePO4 battery pack nominal voltage. 0-40°C operating range covers typical indoor clinical and office environments per the operating environment constraint. | subsystem, drive-subsystem, session-700, idempotency:sub-motor-assembly-power-700
SUB-DRIVESUBSYSTEM-006 | When the Drive Subsystem detects Motor Controller Unit CAN heartbeat absence exceeding 100 ms, the Drive Subsystem SHALL command both motor drivers to coast-to-stop state and assert MCU_FAULT to the Safety Monitor Processor within 50 ms of timeout detection. | subsystem-requirements | Drive Subsystem |  | Test | 71 |  | Motor Controller Unit (hex D4F57A18) is System-Essential; undetected MCU loss leads to silent loss of propulsion authority. 100 ms heartbeat timeout provides two missed 50 Hz frames before declaring fault, balancing false-positive avoidance with response timeliness. | subsystem, drive-subsystem, session-705, sil-2, idempotency:sub-drive-mcu-heartbeat-705
SUB-DRIVESUBSYSTEM-007 | The Motor Controller Unit SHALL implement hardware overcurrent protection independent of the Main Application Processor that disconnects motor phase outputs within 5ms if phase current exceeds 120% of rated maximum, with no firmware execution required for this protection to operate. | subsystem-requirements | Drive Subsystem |  | Test | 86 |  | Motor Controller Unit is classified System-Essential (hex D4F57A18). Firmware-dependent overcurrent protection is insufficient for a SIL-2 drive function: if the MAP is faulted or busy, motor runaway or coil burn is possible within milliseconds. Hardware-only protection eliminates this failure mode and is standard practice in IEC 61800-5-2 (Safety requirements for adjustable speed electrical power drive systems). | subsystem, drive-subsystem, session-706, sil-2, idempotency:sub-drive-mcu-hw-overcurrent-706
SUB-DRIVESUBSYSTEM-008 | While in Degraded/Assisted mode, the Drive Subsystem SHALL enforce a maximum speed cap of 3 km/h by limiting PWM duty cycle to 50% of the Normal Navigation maximum, with the speed cap applied in firmware prior to the Motor Controller Unit command, independent of any BCI or joystick input value. | subsystem-requirements | Drive Subsystem |  | Test | 86 |  | SYS-REQ-022 requires 3 km/h maximum in Degraded mode. Implementing the cap in the Drive Subsystem firmware (not in the BCI pipeline) ensures it applies regardless of BCI or joystick source, preventing the cap from being bypassed if the BCI subsystem transitions incorrectly. | subsystem, drive-subsystem, session-706, sil-2, idempotency:sub-drive-degraded-speed-cap-706
SUB-DRIVESUBSYSTEM-009 | While operating in Care Attendant Override mode, the Drive Subsystem SHALL respond to joystick steering and throttle inputs within 100ms, maintain obstacle detection response within 200ms, limit maximum speed to 6 km/h, and sustain override operation for the full battery duration specified by SUB-REQ-014 without consuming more than 15% additional power versus Normal Navigation mode at equivalent speed. | subsystem-requirements | Drive Subsystem |  | Test | 86 |  | STK-REQ-004 mandates carer override with obstacle detection active but specifies no measurable performance floor for the override mode. 100ms joystick response derived from SYS-REQ-009 (handover latency). 6 km/h matches Normal Navigation maximum. The original 30-minute duration floor at 50% speed was flagged rt-implausible-value because it is far weaker than the 4-hour battery capacity guarantee in SUB-REQ-014, creating an artificial and non-binding constraint. The revised requirement instead bounds override-mode power consumption relative to Normal Navigation, preventing undisclosed high-drain override modes from undermining the SUB-REQ-014 runtime guarantee. | subsystem, drive, session-708, qc, idempotency:sub-drive-override-mode-perf-708, red-team-session-715, reqs-eng-session-717
SUB-DRIVESUBSYSTEM-010 | The Drive Subsystem, including the Motor Controller Unit and Motor Power Isolation Relay, SHALL comply with IEC 60601-1 (Medical electrical equipment — General requirements for basic safety and essential performance) for protection against electrical hazards, and all motor drive electronics SHALL meet the EMC requirements of IEC 60601-1-2 (Electromagnetic compatibility — Requirements and tests for medical electrical equipment). | subsystem-requirements | Drive Subsystem |  | Inspection | 86 |  | Drive subsystem contains 48V DC power electronics in direct proximity to a patient-class user. IEC 60601-1 applies because the EEG Accessibility Buggy is a medical device (confirmed by STK-REQ-010 requirement for design history file and clinical evaluation report). The Motor Controller Unit (MCU) and Motor Power Isolation Relay are classified as Regulated by UHT Substrate but no compliance requirement existed at subsystem level. IEC 60601-1-2 EMC compliance is separately required because the drive PWM switching may interfere with EEG acquisition — this requirement closes that gap with a regulatory floor. | subsystem, drive, compliance, session-708, qc, idempotency:sub-drive-iec60601-compliance-708
SUB-HMISUBSYSTEM-001 | The Audio Alert Module SHALL produce an audible tone of minimum 80dB SPL at 1m when an emergency stop is activated, within 100ms of the E-stop signal being asserted, and SHALL sustain this tone until the system transitions out of the E-stop state. | subsystem-requirements | HMI Subsystem |  | Test | 86 |  | MoP basis: IEC 60601-1-8 (Medical electrical equipment — Alarm systems) clause 6.3.2 specifies a minimum of 65 dBA for high-priority alarm signals in medical environments. The 80dB SPL floor in this requirement adds a 15dB safety margin above the IEC 60601-1-8 baseline to account for clinical environment background noise (hospital ward ambient noise typically 60–70 dB, per HSE noise guidance for NHS environments). STK-REQ-005 sets the stakeholder-level threshold at ≥75 dBA at 1m; SUB-REQ-032 allocates the Audio Alert Module a 5dB margin above the stakeholder minimum, accounting for measurement uncertainty and acoustic absorption at distance. 100ms onset ensures auditory alert is concurrent with physical braking deceleration at 6 km/h (vehicle travels <17cm before alert fires). | subsystem, hmi-subsystem, session-701, idempotency:sub-hmi-audio-estop-701, reqs-eng-session-716
SUB-HMISUBSYSTEM-002 | The Status LED Array SHALL display a distinct colour state for each of: Normal Navigation (green), Degraded/Reduced Mode (amber), E-stop Active (red), Charging (blue), and BCI Calibration (pulsing white), with state transitions completing within 200ms of the triggering event. | subsystem-requirements | HMI Subsystem |  | Demonstration | 86 |  | STK-REQ-005 requires distinct visual alerts for the care attendant. Using the five colours above maps directly to the ConOps operating modes and allows unambiguous state identification without requiring the attendant to read text on the display. 200ms transition time is fast enough to appear instantaneous to the human observer. | subsystem, hmi-subsystem, session-701, idempotency:sub-hmi-led-states-701
SUB-HMISUBSYSTEM-003 | When the care attendant activates the override switch, the HMI Subsystem SHALL transfer joystick command authority to the Drive Subsystem within 100ms, suppressing all BCI command outputs for the duration of override mode. | subsystem-requirements | HMI Subsystem |  | Test | 100 |  | SYS-REQ-009 mandates 100ms joystick authority transfer. This decomposes the latency budget to the HMI/Drive interface. Failure to meet this threshold risks momentary loss of directional authority during handover, which is a collision risk in tight corridors. The 100ms bound is derived from vehicle control latency studies in powered wheelchair standards (ISO 7176-11). | subsystem, hmi-subsystem, drive-subsystem, session-702, idempotency:sub-hmi-joystick-override-handover-702
SUB-HMISUBSYSTEM-004 | The Display Unit SHALL render BCI classification accuracy, vehicle speed, battery state of charge, and active operating mode on a sunlight-readable screen with minimum 300 cd/m2 brightness, updating at minimum 2Hz in Normal Navigation mode. | subsystem-requirements | HMI Subsystem |  | Test | 100 |  | Derived from SYS-REQ-008 (accuracy threshold) and SYS-REQ-018 (calibration readiness). Care attendants require real-time visibility of BCI performance to intervene before accuracy degrades below the 70% transition threshold. A 300 cd/m2 brightness floor ensures readability in hospital dayroom and outdoor settings. 2Hz update prevents perceived lag in a medical device. | subsystem, hmi-subsystem, session-703, idempotency:sub-hmi-display-703
SUB-HMISUBSYSTEM-005 | The Status LED Array SHALL be visible from any angle within a 120-degree horizontal arc to the rear of the vehicle, with a minimum luminous intensity of 10 mcd per LED element, to allow care attendants approaching from behind to read vehicle state. | subsystem-requirements | HMI Subsystem |  | Test | 86 |  | Care attendants typically approach from behind or to the side of the buggy. An angular field of 120 degrees to the rear ensures the attendant can read the vehicle state (green/amber/red) at normal approach angles without requiring a direct line of sight from the front. The 10mcd intensity is derived from ISO 7176-22 (powered wheelchair signal lamp requirements). | subsystem, hmi-subsystem, session-703, idempotency:sub-hmi-led-visibility-703
SUB-HMISUBSYSTEM-006 | When the care attendant activates the rear-mounted override switch, the HMI Subsystem SHALL assert a hardwired CARER_OVERRIDE signal to the Safety Monitor Processor within 50 ms and maintain the signal for the duration of physical switch engagement. | subsystem-requirements | HMI Subsystem |  | Test | 86 |  | SYS-REQ-009 requires care attendant steering transfer within 100 ms end-to-end; the HMI Subsystem accounts for the first 50 ms of that budget through hardwired (not software) assertion to prevent software-layer latency from delaying the takeover. | subsystem, hmi-subsystem, session-705, idempotency:sub-hmi-carer-override-705
SUB-HMISUBSYSTEM-007 | When CARER_OVERRIDE is asserted, the HMI Subsystem SHALL route joystick control commands directly to the Drive Subsystem motor controller at a 50 Hz command rate with latency not exceeding 20 ms from joystick input to CAN frame transmission. | subsystem-requirements | HMI Subsystem |  | Test | 71 |  | SYS-REQ-009 requires joystick steering transfer within 100 ms; 20 ms HMI command latency plus 50 ms safety assertion budget leaves 30 ms for drive subsystem response, satisfying the end-to-end bound at 50 Hz control rate. | subsystem, hmi-subsystem, session-705, idempotency:sub-hmi-joystick-routing-705
SUB-HMISUBSYSTEM-008 | When the care attendant activates the rear-mounted override switch, the HMI Subsystem SHALL complete authority transfer from BCI commands to joystick steering and throttle within 100ms of switch actuation, measured from the falling edge of the CARER_OVERRIDE signal to the first joystick command accepted by the Drive Subsystem. | subsystem-requirements | HMI Subsystem |  | Test | 86 |  | SYS-REQ-009 requires authority transfer within 100ms. This timing is derived from a 6 km/h maximum speed giving 167mm travel per 100ms; exceeding this window allows the vehicle to travel approximately one chair-width during uncontrolled motion, which is unacceptable for a safety-critical handover. The 100ms budget is decomposed as: 20ms switch debounce, 30ms CARER_OVERRIDE propagation, 50ms Drive Subsystem command accept latency. | subsystem, hmi-subsystem, session-706, sil-2, idempotency:sub-hmi-joystick-100ms-706
SUB-HMISUBSYSTEM-009 | When the rolling 2-minute BCI classification accuracy drops below 70%, the Audio Alert Module SHALL emit a 85 dBSPL (at 1m) pulsed tone at 880Hz within 200ms of the accuracy threshold crossing, sustained until the mode transitions or care attendant acknowledges via the HMI. | subsystem-requirements | HMI Subsystem |  | Test | 86 |  | Derives from SYS-REQ-008 which requires audible alert to care attendant when entering Degraded/Assisted mode. 85 dBSPL at 880Hz is audible over typical care facility ambient noise (60-70 dB) without exceeding care setting guidelines. 200ms alert latency is well within the 2-minute rolling window that triggers the mode change, ensuring the care attendant is notified before the buggy is in degraded state for more than a few seconds. | subsystem, hmi-subsystem, sil-0, session-707, idempotency:sub-hmi-audio-care-attendant-707
SUB-HMISUBSYSTEM-010 | When the rear-mounted override switch is activated, the HMI Subsystem SHALL transfer steering and throttle authority to the rear joystick, disable BCI movement command processing, and confirm handover to the Drive Subsystem within 100ms of the override switch activation signal. | subsystem-requirements | HMI Subsystem |  | Test | 86 |  | Derives from SYS-REQ-009 which requires joystick handover within 100ms of care attendant override. The 100ms budget is split: 20ms for switch debounce and signal propagation, 50ms for authority transfer to Drive Subsystem, 30ms for BCI command disable confirmation. This ensures care attendants can take safe control rapidly during user distress or post-degradation handover without perceptible delay. | subsystem, hmi-subsystem, sil-2, session-707, idempotency:sub-hmi-joystick-override-707
SUB-HMISUBSYSTEM-011 | While the system is in Degraded/Assisted mode, the Status LED Array SHALL display a continuous amber pattern at 2Hz pulse rate and the Display Unit SHALL present a persistent text notification reading 'DEGRADED MODE — BCI ACCURACY LOW — SPEED LIMITED TO 3 KM/H' on the primary status screen within 100ms of mode entry, remaining visible until mode change. | subsystem-requirements | HMI Subsystem |  | Demonstration | 86 |  | Derives from SYS-REQ-022 which requires amber LED and text notification on entry to Degraded/Assisted mode. The 2Hz amber pulse is distinct from normal mode (solid green) and emergency stop (1Hz red), providing unambiguous visual status across the three primary operating modes. The text notification includes the speed limit to inform both the user and any bystanders of reduced capability. 100ms display latency is within the HMI response budget after mode transition. | subsystem, hmi-subsystem, sil-0, session-707, idempotency:sub-hmi-degraded-led-display-707
SUB-PERCEPTIONSUBSYSTEM-001 | The Forward Depth Sensor Array SHALL measure distances in the forward 120-degree arc with a range of 0.05m to 2.0m and an accuracy of plus or minus 50mm at no less than 10Hz, under all operating lighting conditions from 0 lux to 10,000 lux. | subsystem-requirements | Perception Subsystem |  | Test | 86 |  | SYS-REQ-006 mandates detection of obstacles within 2m forward zone. Three VL53L5CX sensors in a 120° arc provide full forward coverage with 8x8 depth grids. Accuracy of ±50mm is sufficient to trigger an alert with >1 update cycle margin at 6 km/h (vehicle travels 167mm per 100ms). 10Hz update rate matches minimum safety reaction loop. | subsystem, perception-subsystem, sil-2, session-701, idempotency:sub-perception-fwd-detection-701, red-team-session-715, reqs-eng-session-717
SUB-PERCEPTIONSUBSYSTEM-002 | The Perception MCU SHALL process all sensor inputs and emit an obstacle alert frame to the Safety Monitor Processor within 50ms of sensor data arrival, sustaining this throughput at full 10Hz sensor update rate. | subsystem-requirements | Perception Subsystem |  | Test | 100 |  | At 6 km/h the vehicle travels 83mm in 50ms. Combined with forward sensor range of 2m, the 50ms processing budget leaves a minimum 1.9m braking margin for the Safety Subsystem to act. This derives from SYS-REQ-006 and the 20ms relay cutoff time in SYS-REQ-002. | subsystem, perception-subsystem, sil-2, session-701, idempotency:sub-perception-mcu-latency-701, red-team-session-715, reqs-eng-session-717
SUB-PERCEPTIONSUBSYSTEM-003 | The Side Proximity Sensor Pair SHALL detect objects within 0.5m laterally on each side of the vehicle and deliver a TTL-level alert signal to the Perception MCU within 30ms of intrusion. | subsystem-requirements | Perception Subsystem |  | Test | 100 |  | At 6 km/h during a tight turn, lateral obstacles present risk of entrapment or collision with door frames. The 0.5m detection threshold at 30ms latency provides 83mm travel margin before the safety chain can respond. Derived from STK-REQ-014 (safe stopping distances) and the maximum vehicle width of 700mm per SYS-REQ-010. | subsystem, perception-subsystem, sil-2, session-701, idempotency:sub-perception-side-sensors-701, red-team-session-715, reqs-eng-session-717
SUB-PERCEPTIONSUBSYSTEM-004 | When the Perception MCU fails to produce an obstacle alert frame within 150ms of its scheduled transmission window, the Safety Monitor Processor SHALL treat the timeout as an obstacle-present condition and initiate the emergency stop sequence. | subsystem-requirements | Perception Subsystem |  | Test | 86 |  | IEC 61508 (Functional safety of E/E/PE safety-related systems) SIL 2 requires fail-safe behaviour on subsystem communication loss. Absence of sensor data is a credible failure mode (MCU lockup, I2C bus fault). Treating silence as an obstacle-present condition is the most conservative safe state and prevents motion through undetected obstacles. | subsystem, perception-subsystem, sil-2, safety, session-701, idempotency:sub-perception-safe-state-701
SUB-PERCEPTIONSUBSYSTEM-005 | While vehicle tilt measured by the Inclinometer Tilt Sensor Unit exceeds 15 degrees on any axis, the Perception Subsystem SHALL assert a tilt-hazard signal to the Safety Monitor Processor within 50ms of threshold crossing. | subsystem-requirements | Perception Subsystem |  | Test | 100 |  | SYS-REQ-011 requires braking on 15-degree tilt detection. This decomposes the sensing element to the Perception Subsystem. The 50ms assertion time leaves 150ms for the Safety Monitor Processor to command brake application within the 200ms total SYS-REQ-002 safety response budget. The inclinometer is physically in the Perception subsystem's sensor suite (Electronics Bay mounting). | subsystem, perception-subsystem, sil-2, session-702, idempotency:sub-perception-inclinometer-tilt-702, red-team-session-715, reqs-eng-session-717
SUB-PERCEPTIONSUBSYSTEM-006 | The Inclinometer Tilt Sensor Unit SHALL operate from a regulated 3.3V supply derived from the DC-DC Converter Array, with a maximum continuous current draw of 10mA, and SHALL maintain measurement accuracy within specification for supply voltage variation of 3.0V to 3.6V. | subsystem-requirements | Perception Subsystem |  | Test | 86 |  | Lint finding: inclinometer is Powered but has no power source requirements. The inclinometer is a MEMS device (e.g. ST LSM6DSO) operating at 3.3V logic level consistent with the Perception MCU. Defining the power envelope prevents sensor brown-out during load transients from the drive motors, which would cause spurious tilt alarms and emergency stops. | subsystem, perception-subsystem, power, sil-2, session-702, idempotency:sub-perception-inclinometer-power-702, red-team-session-715, reqs-eng-session-717
SUB-PERCEPTIONSUBSYSTEM-007 | While vehicle tilt measured by the onboard inclinometer exceeds 15 degrees in any axis, the Perception Subsystem SHALL continuously publish tilt angle, axis identity, and confidence score to the Safety Monitor Processor at no less than 20 Hz. | subsystem-requirements | Perception Subsystem |  | Test | 86 |  | SYS-REQ-011 triggers automated halt on 15-degree tilt; the Perception Subsystem must provide tilt data at sufficient rate (20 Hz) to ensure the Safety Subsystem can detect and respond within 200 ms before the system enters an unsafe attitude. | subsystem, perception-subsystem, session-705, idempotency:sub-perception-inclinometer-705, red-team-session-715, reqs-eng-session-717
SUB-PERCEPTIONSUBSYSTEM-008 | When one or more elements of the Forward Depth Sensor Array fail to deliver a valid distance frame within 200ms, the Perception MCU SHALL reduce commanded maximum vehicle speed to 1.5 km/h, assert a sensor-fault alert to the Safety Monitor Processor, and maintain this reduced-speed safe state until all sensor elements are confirmed operational. | subsystem-requirements | Perception Subsystem |  | Test |  |  | Partial failure of the forward sensor array reduces obstacle coverage below the full 120-degree arc, creating blind zones. At 1.5 km/h, the vehicle travels 83mm per 200ms safety response cycle, maintaining a minimum 1.9m stopping margin within the 2m sensor range. This degraded safe state allows continued slow-speed operation; total perception failure escalates to emergency stop per the Perception MCU watchdog requirement. | subsystem, perception-subsystem, sil-2, session-717, reqs-eng-session-717, failure-mode, safe-state, addresses:SUB-REQ-023, idempotency:fm-sub-023-fwd-sensor-dropout-717
SUB-PERCEPTIONSUBSYSTEM-009 | When the Perception MCU fails to emit an obstacle-alert frame to the Safety Monitor Processor for more than 150ms, the Safety Monitor Processor SHALL assert a perception-fault condition: engage emergency braking, disable all drive motor commands, and inhibit restart until the Perception MCU heartbeat is confirmed operational. | subsystem-requirements | Perception Subsystem |  | Test |  |  | At 6 km/h, a 150ms perception blackout allows 250mm of undetected travel within the 2m sensor range. The Safety Monitor Processor must treat heartbeat loss as equivalent to an obstacle detection event. 150ms timeout is 1.5 times the nominal 100ms sensor-MCU processing cycle, allowing one missed frame before fault assertion. Derived from SYS-REQ-002 (200ms total safety response budget) and SYS-REQ-006 (obstacle detection). | subsystem, perception-subsystem, sil-2, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-024, idempotency:fm-sub-024-perception-mcu-watchdog-717
SUB-PERCEPTIONSUBSYSTEM-010 | When either Side Proximity Sensor element fails to respond to its health-check poll within 200ms, the Perception MCU SHALL assert a side-sensor-fault flag to the Safety Monitor Processor and restrict maximum vehicle speed to 0.5 km/h until the fault is cleared by an authorised operator or maintenance reset. | subsystem-requirements | Perception Subsystem |  | Test |  |  | Side proximity sensor failure creates a blind zone for lateral obstacles during turns. At 6 km/h in a 700mm-wide vehicle in 900mm doorways this presents entrapment risk per the Daily Independent Navigation ConOps scenario. Reducing speed to 0.5 km/h provides minimum 100ms carer reaction margin to apply emergency stop before a lateral collision. Derived from STK-REQ-014 (safe stopping distances) and doorway clearance analysis. | subsystem, perception-subsystem, sil-2, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-025, idempotency:fm-sub-025-side-sensor-failure-717
SUB-PERCEPTIONSUBSYSTEM-011 | When the Inclinometer Tilt Sensor Unit fails to deliver a valid tilt measurement to the Perception MCU within 200ms (sensor timeout or invalid checksum on three consecutive frames), the Safety Monitor Processor SHALL treat the failure as a tilt-threshold-exceeded condition and command emergency stop, maintaining halt until the inclinometer is confirmed operational. | subsystem-requirements | Perception Subsystem |  | Test |  |  | An inclinometer that goes silent may have failed at or beyond the 15-degree threshold, creating an undetected tip-over hazard. Conservative fail-safe policy per IEC 61508 (Functional safety of E/E/PE safety-related systems) SIL-2 fault-detection requirements. 200ms timeout is 4 times the nominal 50ms assertion time, providing three missed-update margin before fault declaration to avoid nuisance stops. | subsystem, perception-subsystem, sil-2, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-035, idempotency:fm-sub-035-tilt-sensor-silence-717
SUB-PERCEPTIONSUBSYSTEM-012 | When the Perception Subsystem fails to publish valid tilt data to the Safety Monitor Processor for more than 100ms during active navigation, the Safety Monitor Processor SHALL treat the publication silence as a tilt-threshold-exceeded condition and command emergency stop, maintaining halt until tilt data publication is confirmed restored. | subsystem-requirements | Perception Subsystem |  | Test |  |  | Tilt-data absence during navigation is treated as tilt-threshold exceedance per fail-safe policy. At 20Hz publication rate, 100ms silence represents two missed updates — sufficient margin to distinguish transient jitter from sensor failure without nuisance stops. This ensures the 15-degree tilt halt condition in SYS-REQ-011 cannot be bypassed by a publication failure in the Perception Subsystem. | subsystem, perception-subsystem, sil-2, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-054, idempotency:fm-sub-054-tilt-pubfailure-717
SUB-POWERSUBSYSTEM-001 | The Battery Management System SHALL disconnect the battery pack from the 48V bus within 200ms when any cell temperature exceeds 60°C or any cell voltage exceeds 3.65V. | subsystem-requirements | Power Subsystem |  | Test | 100 |  | Directly implements SYS-REQ-012 at the BMS level. 200ms disconnect time is the IEC 62133 (Secondary lithium cells — Safety requirements) maximum allowable time before sustained OV or OT can initiate thermal runaway in LiFePO4 chemistry. Failure to disconnect risks irreversible thermal runaway with fire hazard near a vulnerable user. | subsystem, power-subsystem, sil-3, session-699, idempotency:sub-pwr-bms-thermal-cutoff-699, red-team-session-715, reqs-eng-session-717
SUB-POWERSUBSYSTEM-002 | The Lithium Iron Phosphate Battery Pack SHALL deliver a minimum of 4 hours continuous operation at rated system load (drive motors, BCI processing, HMI, safety systems combined) before state of charge falls below 10%. | subsystem-requirements | Power Subsystem |  | Test | 100 |  | Directly implements SYS-REQ-004 at subsystem level. 10% SoC floor preserves headroom for safe-stop execution and BMS shutdown sequence without total power loss. 4-hour operational duration is the STK-derived minimum for full-day facility use between charges. SIL-3 classification rationale: the 10% SoC floor is the trigger condition for the SIL-3 safe-stop sequence defined in the companion failure-mode requirement (idempotency:fm-sub-014-battery-socdepleted-717). The battery pack capacity requirement is the performance prerequisite for that safety function — battery depletion before 10% SoC remaining would prevent safe-stop execution, resulting in uncontrolled motion hazard H-001 and tip-over hazard H-005. SIL-3 is appropriate because loss of this function could result in serious or fatal injury to the EEG-impaired user. | subsystem, power-subsystem, session-699, idempotency:sub-pwr-battery-runtime-699, red-team-session-715, reqs-eng-session-717
SUB-POWERSUBSYSTEM-003 | The DC-DC Converter Array SHALL maintain each output rail (12V, 5V, 3.3V) within ±5% of nominal voltage under all load conditions from 10% to 100% rated current, with output ripple not exceeding 50mV peak-to-peak. | subsystem-requirements | Power Subsystem |  | Test | 100 |  | IEC 60601-1 (Medical electrical equipment — General requirements for basic safety) requires stable power supply to safety-critical electronics. ±5% regulation covers the input tolerance range of embedded processors and wireless modules. 50mV ripple is the maximum specification for the BCI processing modules' switching noise sensitivity on EEG acquisition circuitry. | subsystem, power-subsystem, sil-3, session-699, idempotency:sub-pwr-dcdc-regulation-699, red-team-session-715, reqs-eng-session-717
SUB-POWERSUBSYSTEM-004 | The Charge Controller SHALL charge the battery pack from 20% to 100% state of charge within 3 hours when connected to a 240V AC facility supply. | subsystem-requirements | Power Subsystem |  | Test | 100 |  | SYS-REQ-014 requirement cascaded directly. 3-hour charge time allows overnight charging on standard facility schedules (facility changeover at 18:00, vehicle available by 21:00 for next-day use). Faster charging is not feasible within the 3kW facility outlet rating while maintaining the CC/CV profile required for LiFePO4 longevity. SIL escalation resolution: the SIL-3 tag on this requirement reflects its role as the primary charge timing specification for the SIL-3 power system. However, the charge timing function itself is not the SIL-3 safety function — the mains fault isolation function is (addressed in companion requirement idempotency:fm-sub-016-charger-mainsfault-717 at SIL-2). This requirement's SIL-3 tag should be understood as marking it as part of the SIL-3 power subsystem scope, not as claiming the charge timing is itself a SIL-3 safety function. The IEC 62368-1 fault isolation companion requirement carries the appropriate SIL-2 classification for the fault-isolation function. | subsystem, power-subsystem, session-699, idempotency:sub-pwr-charge-time-699, red-team-session-715, reqs-eng-session-717
SUB-POWERSUBSYSTEM-005 | The Charge Controller SHALL accept charging input from a facility charging dock providing 230V 50Hz or 120V 60Hz AC mains, and SHALL complete pack charge from 20% to 100% state of charge within 4 hours at each nominal supply voltage. | subsystem-requirements | Power Subsystem |  | Test | 86 |  | SYS-REQ-014 specifies 4-hour charging from facility dock at both EU and US mains voltages. The Charge Controller is the Power Subsystem component responsible for dock interface and CC/CV charge management. Compliance with both supply standards is required for international deployment in care facilities. | subsystem, power-subsystem, session-702, idempotency:sub-power-charge-controller-facility-dock-702, red-team-session-715, reqs-eng-session-717
SUB-POWERSUBSYSTEM-006 | When connected to a facility charging dock providing 230V 50Hz AC supply, the Power Subsystem onboard charger SHALL limit inrush current to less than 16A peak, regulate charging current to maintain cell temperature below 45°C, and complete charge from 20% to 100% SoC within 4 hours. | subsystem-requirements | Power Subsystem |  | Test | 86 |  | SYS-REQ-014 sets the 4-hour charge window; the Power Subsystem must manage inrush and thermal limits to use facility power safely and meet the charge time without exceeding the 16 A standard socket circuit protection threshold. | subsystem, power-subsystem, session-705, idempotency:sub-power-facility-charging-705, red-team-session-715, reqs-eng-session-717, tech-author-session-719
SUB-POWERSUBSYSTEM-007 | The Charge Controller SHALL accept 230V 50Hz or 120V 60Hz mains input via the facility charging dock connector and charge the Lithium Iron Phosphate Battery Pack from 20% to 100% State of Charge within 4 hours using CC-CV charging profile, with automatic termination and trickle maintenance. | subsystem-requirements | Power Subsystem |  | Test | 86 |  | Derives from SYS-REQ-014 which specifies the 4-hour charge window from 20% to 100% SOC. The LiFePO4 48V 30Ah pack requires approximately 28.8Ah at 48V = 1.38 kWh from 20% to 100%. A CC-CV profile at 10A charge current achieves this in under 3.5 hours with end-of-charge cut-off, providing margin against cell aging degradation. Dual-voltage acceptance (230V/120V) enables international facility deployment. | subsystem, power-subsystem, sil-1, session-707, idempotency:sub-power-charge-controller-dock-707, red-team-session-715, reqs-eng-session-717
SUB-POWERSUBSYSTEM-008 | When the Battery Management System fails to disconnect the battery pack from the 48V bus within 250ms of an overvoltage or overtemperature detection, the Safety Monitor Processor SHALL detect the BMS watchdog timeout and assert system-wide emergency stop within 50ms. | subsystem-requirements | Power Subsystem |  | Test |  |  | BMS disconnect failure leaves LiFePO4 cells at sustained overvoltage or overtemperature, the initiating conditions for thermal runaway per IEC 62133 (Secondary lithium cells — Safety requirements). The Safety Monitor Processor's independent watchdog on the BMS status line ensures the safety chain remains active even if the BMS processor itself fails. 250ms detection covers the 200ms BMS disconnect budget plus 50ms measurement margin. | subsystem, power-subsystem, sil-3, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-013, idempotency:fm-sub-013-bms-disconnect-717
SUB-POWERSUBSYSTEM-009 | When the Lithium Iron Phosphate Battery Pack state of charge reaches 10%, the Power Subsystem SHALL activate a safe-stop sequence: halt all drive motor commands within 5 seconds, engage electromechanical parking brakes, sustain Safety Monitor Processor and emergency braking power for a minimum of 5 minutes, and assert a low-battery HMI alert. | subsystem-requirements | Power Subsystem |  | Test |  |  | Loss of drive power at SoC below 10% without a controlled safe-stop creates H-005 (tip-over) and H-001 (uncontrolled motion) hazards if brakes are not engaged before total power loss. Five-minute emergency power reserve covers safe user evacuation per STK-REQ-002 minimum carer response time. This failure-mode requirement also addresses the rt-sil-escalation finding on SUB-REQ-014: the SIL-3 classification applies to the safety-power maintenance function, not to the 4-hour runtime performance floor. | subsystem, power-subsystem, sil-3, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-014, idempotency:fm-sub-014-battery-socdepleted-717
SUB-POWERSUBSYSTEM-010 | When any DC-DC Converter Array output rail deviates more than 10% from nominal voltage for more than 100ms, the Safety Monitor Processor SHALL assert a power-fault condition, command emergency braking, disable all drive motor control signals, and maintain power-fault safe state until the rail recovers to within 5% of nominal and an authorised operator clears the fault. | subsystem-requirements | Power Subsystem |  | Test |  |  | Processor and sensor power deviation beyond 10% causes undefined behaviour in BCI processing, Perception MCU, and Safety Monitor Processor sub-circuits. Transition to power-fault safe state (drives off, brakes on) prevents uncontrolled motion during supply instability per IEC 60601-1 (Medical electrical equipment — General requirements for basic safety). The 100ms tolerance window matches the minimum hold-up time of DC-DC filter capacitors before output degrades below logic-level thresholds. | subsystem, power-subsystem, sil-3, session-717, reqs-eng-session-717, failure-mode, safe-state, addresses:SUB-REQ-015, idempotency:fm-sub-015-dcdc-railfault-717
SUB-POWERSUBSYSTEM-011 | When the Charge Controller detects a mains supply fault (supply voltage outside 207V-253V AC, frequency outside 47Hz-53Hz, or ground fault current exceeding 10mA RMS), it SHALL disconnect from the facility mains supply within 200ms, lock the charge port relay, and assert a charge-fault alert on the HMI display. | subsystem-requirements | Power Subsystem |  | Test |  |  | Mains fault conditions during charging present fire and electric shock hazards in a care environment. 200ms disconnect is within IEC 62368-1 (Audio/video and IT equipment — Safety requirements) fast overcurrent protection requirements. Ground fault detection at 10mA RMS is required by IEC 60364-4-41 for care environments where users have reduced protective reflexes. This also addresses the rt-sil-escalation finding on SUB-REQ-016: the SIL-relevant function is fault isolation, not the 3-hour charge timing. | subsystem, power-subsystem, sil-2, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-016, idempotency:fm-sub-016-charger-mainsfault-717
SUB-POWERSUBSYSTEM-012 | When the Charge Controller detects mains supply voltage outside 207V-253V AC at 50Hz or 102V-132V AC at 60Hz at the facility dock connector, it SHALL reject the connection, maintain galvanic isolation from the supply, and display a supply-incompatible warning on the HMI. | subsystem-requirements | Power Subsystem |  | Test |  |  | Operating the Charge Controller on an out-of-range supply risks uncontrolled charging current that could trigger thermal runaway in the LiFePO4 pack. Maintaining galvanic isolation as the rejection safe state aligns with IEC 62368-1 (Audio/video and IT equipment — Safety requirements) energy hazard reduction requirements. Voltage acceptance windows are plus or minus 10% of nominal per IEC 60038 standard voltages. | subsystem, power-subsystem, sil-2, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-036, idempotency:fm-sub-036-charger-volt-reject-717
SUB-POWERSUBSYSTEM-013 | When the 3.3V supply to the Inclinometer Tilt Sensor Unit deviates outside the 3.0V-3.6V operating range for more than 50ms, the DC-DC Converter Array health monitor SHALL assert a sensor-power fault to the Safety Monitor Processor, causing immediate vehicle halt and inhibiting motion until the 3.3V rail returns to within specification. | subsystem-requirements | Power Subsystem |  | Test |  |  | Inclinometer operation below 3.0V produces out-of-specification tilt readings, creating false-negative tilt alarms and potential for undetected tip-over. 50ms fault assertion tolerance matches the maximum hold-up time of 3.3V rail decoupling capacitors before output degrades below the sensor minimum operating threshold. The safety impact is equivalent to tilt-sensor failure under SYS-REQ-011 and warrants immediate halt. | subsystem, power-subsystem, sil-2, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-040, idempotency:fm-sub-040-inclinometer-rail-fault-717
SUB-POWERSUBSYSTEM-014 | When inrush current to the Power Subsystem charger exceeds 16A peak or cell temperature during charging exceeds 45°C for more than 10 seconds, the Charge Controller SHALL interrupt the charge cycle within 100ms, isolate from the facility mains supply, and assert a charge-fault alert that persists until acknowledged by an authorised operator. | subsystem-requirements | Power Subsystem |  | Test |  |  | Inrush exceeding 16A may trip facility circuit breakers per BS EN 60898-1 (Electrical accessories — Circuit-breakers for overcurrent protection) 16A Type B/C trip curves. Cell temperature exceeding 45 degrees C indicates abnormal electrochemical activity; IEC 62133 (Secondary lithium cells — Safety requirements) specifies 45 degrees C as the maximum permissible cell temperature during charging. Persistent fault state prevents automatic resumption without operator acknowledgement, complying with IEC 60601-1 alarm persistence requirements. | subsystem, power-subsystem, sil-2, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-055, idempotency:fm-sub-055-charge-inrush-thermal-717, tech-author-session-719
SUB-POWERSUBSYSTEM-015 | When the Charge Controller fails to reach 100% state of charge within 5 hours of charge initiation, it SHALL terminate the charge cycle, log a battery-health fault event with timestamp to the CMMS interface, and alert the facility operator via the HMI display. | subsystem-requirements | Power Subsystem |  | Test |  |  | CC-CV charging the 48V 30Ah LiFePO4 pack from 20% to 100% requires approximately 1.4 kWh; at 10A charge current this completes in 3.5 hours under normal conditions. Five hours allows for capacity degradation to approximately 70% of rated before flagging — at which point battery replacement is required to maintain the 4-hour operational duration in SUB-REQ-014. Failure to terminate indicates abnormal internal resistance increase per IEC 62133 (Secondary lithium cells — Safety requirements) and requires maintenance assessment. | subsystem, power-subsystem, sil-1, session-717, reqs-eng-session-717, failure-mode, addresses:SUB-REQ-071, idempotency:fm-sub-071-charge-no-termination-717
SUB-SAFETYSUBSYSTEM-001 | The Safety Monitor Processor SHALL operate on a processor with no shared memory, clock domain, or power rail with the main application processor, and SHALL maintain its safety state machine execution at 200Hz even when the main application processor is in fault state. | subsystem-requirements | Safety Subsystem |  | Test | 86 |  | IEC 61508 (Functional safety of E/E/PE safety-related systems) SIL 3 requires that safety functions be independent of the control path. If the Safety Monitor Processor shared resources with the main processor, a single processor fault could simultaneously disable vehicle control and safety response, violating independence and allowing uncontrolled motion — hazards H-001 and H-006. | subsystem, safety-subsystem, sil-3, session-697, idempotency:sub-safety-smp-independence-697
SUB-SAFETYSUBSYSTEM-002 | The Motor Power Isolation Relay SHALL disconnect the 48V motor drive power rail within 20ms of receiving the de-energise command from the Safety Monitor Processor, under all load conditions from 0A to 200A. | subsystem-requirements | Safety Subsystem |  | Test | 100 |  | The 200ms E-stop response budget allocated in SYS-REQ-002 (derived from H-001 severity and stopping distance analysis at 6 km/h maximum speed) requires relay actuation <20ms to allow margin for brake actuation time. Mechanical relay de-energisation is the single longest latency element; a longer response risks the vehicle travelling an additional 33mm per ms, which at 20ms already represents 33cm of uncontrolled motion. | subsystem, safety-subsystem, sil-3, session-697, idempotency:sub-safety-relay-response-697
SUB-SAFETYSUBSYSTEM-003 | The Seizure Detection Module SHALL detect generalised spike-and-wave EEG activity exceeding 3 Hz amplitude and >30µV peak-to-peak across all acquisition channels and output a seizure flag to the Safety Monitor Processor within 150ms of seizure onset, with a false positive rate not exceeding 1 per 8 hours of operation. | subsystem-requirements | Safety Subsystem |  | Test | 86 |  | H-006 (catastrophic severity) requires seizure detection to precede motor response — 150ms detection plus relay actuation leaves within the 200ms SYS-REQ-002 window. False positive limit of 1 per 8-hour session is derived from stakeholder acceptability analysis: more frequent nuisance stops would cause users to abandon the system, defeating its purpose as a mobility aid. Algorithm validated against CHB-MIT epilepsy EEG dataset. | subsystem, safety-subsystem, sil-3, session-697, idempotency:sub-safety-seizure-detect-697
SUB-SAFETYSUBSYSTEM-004 | The Inclinometer Tilt Sensor Unit SHALL measure vehicle pitch and roll continuously at 100Hz with accuracy of ±0.5° and SHALL trigger an E-stop command via the Safety Monitor Processor when tilt in any axis exceeds 15 degrees for more than 200ms. | subsystem-requirements | Safety Subsystem |  | Test | 86 |  | SYS-REQ-011 derives from H-005 (critical severity): vehicle tipping on slopes where user cannot self-extract. 15° threshold is derived from EN 12184 (electrically powered wheelchairs) stability testing requirements — vehicles must not tip under static conditions up to this angle. 200ms debounce prevents nuisance stops from dynamic bumps while still catching sustained tilt indicating a genuine tip-over condition. | subsystem, safety-subsystem, sil-2, session-697, idempotency:sub-safety-tilt-trigger-697
SUB-SAFETYSUBSYSTEM-005 | The Manual Emergency Stop Button SHALL be hardwired directly to the Motor Power Isolation Relay coil circuit in series, such that button actuation disconnects motor power without software mediation, within 5ms of button contact closure. | subsystem-requirements | Safety Subsystem |  | Test | 100 |  | Software-mediated E-stop cannot satisfy SIL 3 (IEC 61508) because it introduces a single point of failure in the software execution path. Hardwired series circuit means no combination of software fault, processor reset, or firmware bug can prevent E-stop actuation. 5ms latency is achievable with direct relay contact in series — it is the emergency stop mechanism of last resort required by EN 12184 for powered wheelchairs. | subsystem, safety-subsystem, sil-3, session-697, idempotency:sub-safety-estop-hardwire-697
SUB-SAFETYSUBSYSTEM-006 | When any safety trigger is received by the Safety Monitor Processor — including BCI signal loss notification, seizure flag, tilt threshold exceeded, manual E-stop, or main processor heartbeat failure — the Safety Subsystem SHALL transition to and maintain Emergency Stop state within 200ms, including relay de-energisation, brake command, and HMI alert outputs. | subsystem-requirements | Safety Subsystem |  | Test | 86 |  | IEC 61508 SIL 3 requires a defined safe state for every safety function. The safe state for this system is motor power removed and brakes engaged. 200ms total budget is derived from stopping distance analysis: at 6 km/h maximum speed, a 200ms response results in 33cm of forward travel before braking begins — within the 2m obstacle detection zone defined in SYS-REQ-006, leaving 1.67m braking distance for the mechanical brake. | subsystem, safety-subsystem, sil-3, session-697, idempotency:sub-safety-safe-state-697
SUB-SAFETYSUBSYSTEM-007 | The Safety Monitor Processor SHALL monitor a heartbeat signal from the main application processor, and SHALL trigger Emergency Stop if the heartbeat is absent for more than 500ms or if the heartbeat period deviates by more than 20% from the nominal 100ms period. | subsystem-requirements | Safety Subsystem |  | Test | 86 |  | Main processor failure (software hang, voltage fault, thermal fault) is not detectable by the main processor itself. Watchdog heartbeat is the only mechanism by which the independent Safety Monitor Processor can detect this class of failure and transition to safe state. 500ms absence timeout: derived from worst-case main processor boot time (allowing 5 missed beats) versus the need to detect a genuine hang within 2 processor cycles. | subsystem, safety-subsystem, sil-3, session-697, idempotency:sub-safety-watchdog-697
SUB-SAFETYSUBSYSTEM-008 | The Safety Subsystem SHALL be implemented as a physically separate electronics module from the Main Application Processor, housed within the Electronics Bay on an independent PCB with independent power input, to ensure that a hardware fault in the main processing chain cannot corrupt safety function execution. | subsystem-requirements | Safety Subsystem |  | Inspection | 86 |  | Lint finding: safety subsystem lacks Physical Object trait but has physical embodiment constraints. IEC 61508 SIL-3 requires independence between safety function and non-safety processing to prevent common-cause failures. The physical separation on an independent PCB with independent power ensures the safety monitor can operate when the main processor fails, which is the primary failure mode requiring emergency braking. | subsystem, safety-subsystem, sil-3, session-702, idempotency:sub-safety-physical-embodiment-702
SUB-SAFETYSUBSYSTEM-009 | When the facility emergency system transmits a halt command via the 2.4 GHz wireless interface, the Safety Subsystem SHALL apply both motor power isolation relays and engage electromagnetic braking within 150 ms of signal receipt. | subsystem-requirements | Safety Subsystem |  | Test | 100 |  | SYS-REQ-005 mandates facility emergency integration but assigns no subsystem-level implementation. The Safety Subsystem owns motor de-energisation; 150 ms is derived from the 200 ms SYS-level bound minus 50 ms signal propagation budget for the worst-case wireless link. | subsystem, safety-subsystem, session-705, idempotency:sub-safety-facility-esm-705
SUB-SAFETYSUBSYSTEM-010 | When the Safety Monitor Processor detects Main Application Processor watchdog timeout exceeding 200 ms, the Safety Subsystem SHALL assert SAFE_STOP, halt all motor commands, and log a processor fault record to non-volatile storage. | subsystem-requirements | Safety Subsystem |  | Test | 100 |  | Main Application Processor (hex D2F51008) is System-Essential; loss of the MAP without a safety response leaves the drive subsystem without commands, creating an uncontrolled state. IEC 61508 SIL-2 requires defined response to essential processor failure. | subsystem, safety-subsystem, session-705, sil-2, idempotency:sub-safety-map-watchdog-705
SUB-SAFETYSUBSYSTEM-011 | The Safety Subsystem Safety Monitor Processor SHALL be developed and validated in accordance with IEC 61508-3 (Functional Safety of Electrical/Electronic/Programmable Electronic Safety-related Systems — Part 3: Software requirements) at SIL 3, with documented V-model lifecycle artefacts including hazard and risk analysis, software safety requirements specification, and independent verification. | subsystem-requirements | Safety Subsystem |  | Inspection | 86 |  | Safety Monitor Processor (hex D5F37858) is Regulated and Institutionally Defined; IEC 61508 (Functional safety of E/E/PE safety-related systems) is the applicable standard for safety-critical embedded processors in this domain. SIL 3 is inherited from the Safety Subsystem parent requirement hazard classification. | subsystem, safety-subsystem, session-705, sil-3, standards, idempotency:sub-safety-smp-iec61508-705
SUB-SAFETYSUBSYSTEM-012 | The Manual Emergency Stop Button SHALL use a dual-channel NC (normally-closed) contact circuit wired in series with the Motor Power Isolation Relay coil, such that loss of either channel independently de-energises the relay and halts motor drive within 20ms. | subsystem-requirements | Safety Subsystem |  | Test | 86 |  | Emergency stop is classified System-Essential (hex 408D6AC0); a single-channel E-stop circuit is a single point of failure on a SIL-3 function. Dual NC contacts ensure any wiring fault or contact failure still de-energises the relay, meeting IEC 62061 PLd requirements for safety function SF1. | subsystem, safety-subsystem, session-706, sil-3, idempotency:sub-safety-estop-dual-channel-706
SUB-SAFETYSUBSYSTEM-013 | When the Safety Monitor Processor detects Main Application Processor failure (watchdog timeout, CRC fault, or voltage rail collapse), the Safety Subsystem SHALL assume sole command authority, hold current velocity to zero, and maintain Emergency Stop state until manual reset, with no reliance on MAP firmware for this transition. | subsystem-requirements | Safety Subsystem |  | Test | 86 |  | Main Application Processor is classified System-Essential (hex D2F51008). Without explicit failover specification, a MAP crash could leave the drive system in an undefined state. This requirement ensures the Safety Monitor Processor has unconditional authority on MAP failure, consistent with IEC 61508 SIL-3 defensive design. | subsystem, safety-subsystem, session-706, idempotency:sub-safety-map-failover-706, sil-2
SUB-SAFETYSUBSYSTEM-014 | The Safety Subsystem SHALL be certified to Safety Integrity Level 3 (SIL 3) per IEC 61508 (Functional safety of electrical/electronic/programmable electronic safety-related systems), with a third-party functional safety assessment covering hardware fault tolerance, systematic capability, and software development process. | subsystem-requirements | Safety Subsystem |  | Inspection | 86 |  | Safety Subsystem is classified Regulated (hex D7F73058). The EEG buggy is a medical-adjacent personal mobility device for users who cannot remove themselves from a hazardous situation during a seizure event. SIL-3 certification is required because a Safety Subsystem failure combined with a seizure event (probability of demand: multiple times per year for epileptic users) gives an unacceptable combined risk. Third-party assessment is required by the UK Medical Device Regulations 2002 (as amended) for Class IIa active devices. | subsystem, safety-subsystem, session-706, sil-3, idempotency:sub-safety-sil3-cert-706
SUB-SAFETYSUBSYSTEM-015 | The Inclinometer Tilt Sensor Unit SHALL measure vehicle pitch and roll at 20Hz, and when tilt exceeds 12 degrees in any single axis or 10 degrees combined vector, SHALL issue a pre-warning signal to the Safety Monitor Processor; when tilt exceeds 15 degrees in any axis the Safety Monitor Processor SHALL command immediate brake application and motion inhibit within 50ms. | subsystem-requirements | Safety Subsystem |  | Test | 86 |  | Derives from SYS-REQ-011 which requires braking on 15-degree tilt detection. A pre-warning at 12 degrees gives the Safety Monitor Processor 150ms lead time (at typical incline approach speeds of 0.8 km/h) before the 15-degree hard stop, enabling softer brake application and reducing tip-over risk. The 50ms response time from threshold crossing to brake application is necessary to arrest motion before tilt increases to 18 degrees (tip threshold) at maximum buggy speed. | subsystem, safety-subsystem, sil-2, session-707, idempotency:sub-safety-inclinometer-tilt-707
SUB-VEHICLEPLATFORM-001 | The Chassis Frame SHALL withstand a static load of 150kg distributed across the seat mounting points without permanent deformation, and a dynamic impact load of 3g vertical shock without fracture, in accordance with ISO 7176-8 (Static, impact, and fatigue strengths). | subsystem-requirements | Vehicle Platform |  | Test | 86 |  | A 150kg gross vehicle weight covers the 99th percentile adult occupant mass plus battery and electronics. The 3g shock load covers a threshold crossing at 6 km/h. ISO 7176-8 is the applicable standard for powered wheelchair structural integrity; failure to meet this would result in frame collapse under foreseeable use conditions. | subsystem, vehicle-platform, session-701, idempotency:sub-vp-chassis-load-701, red-team-session-715, reqs-eng-session-717
SUB-VEHICLEPLATFORM-002 | The Electronics Bay SHALL maintain an internal temperature below 70°C for all electronics mounted within it while dissipating up to 40W total under an ambient temperature of 40°C, with the vehicle stationary and no forced air flow. | subsystem-requirements | Vehicle Platform |  | Test | 86 |  | The processor and power conversion components within the Electronics Bay have a maximum junction temperature of 85 degrees Celsius; maintaining the bay below 70 degrees provides a 15 degree thermal margin. 40W is the worst-case total dissipation of all hosted components at peak load. Passive cooling must work with the vehicle stationary as active cooling adds failure modes. | subsystem, vehicle-platform, session-701, idempotency:sub-vp-ebay-thermal-701, tech-author-session-719
SUB-VEHICLEPLATFORM-003 | The Seat and Postural Support System SHALL accommodate occupants from 40kg to 120kg body mass and from 1400mm to 1900mm stature, providing a minimum of 15 degrees of backrest recline adjustment and a lap belt that tightens to resist forward displacement under 3g deceleration. | subsystem-requirements | Vehicle Platform |  | Test | 86 |  | ISO 7176-15 (Wheelchairs — Requirements and test methods for wheelchairs used in motor vehicles) defines the occupant range for accessibility equipment. 3g deceleration matches the emergency stop force at 6 km/h. Failure to restrain the occupant during E-stop is a direct patient safety hazard. | subsystem, vehicle-platform, session-701, idempotency:sub-vp-seat-occupant-701
SUB-VEHICLEPLATFORM-004 | The Wheel and Caster Assembly SHALL provide a turning radius not exceeding 600mm (measured from vehicle centre to outer wheel track) to allow navigation through standard 900mm accessible doorways with a minimum 150mm clearance on each side, while maintaining traction on surfaces inclined up to 6 degrees. | subsystem-requirements | Vehicle Platform |  | Test | 86 |  | Derived from SYS-REQ-010 (700mm vehicle footprint and accessibility environments). A 900mm doorway minus 150mm clearance each side leaves 600mm for the vehicle width plus the swept turning radius. Compliance with BS 8300 (Design of an accessible and inclusive built environment) requires navigation through standard accessible doorway widths. 6-degree gradient is the maximum ramp slope in UK building regulations Part M. | subsystem, vehicle-platform, session-703, idempotency:sub-vp-wheel-turning-703
SUB-VEHICLEPLATFORM-005 | The Electronics Bay SHALL provide IP54 environmental protection per IEC 60529 (IEC 60529: protected against dust ingress at IP5X — wire probe 1mm dia. does not penetrate; protected against water spray from any direction at IP4X), maintaining this rating through all maintenance access cycles, with a minimum 100 maintenance access cycles before seal degradation below IP54. | subsystem-requirements | Vehicle Platform |  | Test | 86 |  | Derived from SYS-REQ-010 and expected deployment in care facilities. Hospital environments include wet floor cleaning with mop spray and occasional outdoor transport. IP54 is the minimum medical device enclosure rating for mobile use (IEC 60601-1 clause 11.6.5). The 100-cycle maintenance life ensures IP rating persists through the 5-year expected service life at quarterly inspection intervals. | subsystem, vehicle-platform, session-703, idempotency:sub-vp-electronics-ip54-703
SUB-VEHICLEPLATFORM-006 | The Vehicle Platform Electronics Bay SHALL provide a USB-C 3.1 Gen 1 service port that exposes an automated diagnostic test suite capable of verifying motor response (peak torque within 5% of nominal), brake torque (hold force >= 80N), sensor calibration (all axes within 2% of reference), battery cell health (internal resistance and capacity), and BCI pipeline integrity (signal-to-noise ratio >= 20dB), with full suite completion within 20 minutes. | subsystem-requirements | Vehicle Platform |  | Demonstration | 86 |  | Derives from SYS-REQ-015 which requires an automated diagnostic suite via USB-C service port. USB-C 3.1 Gen 1 provides 5 Gbps bandwidth and 5A power delivery for bench test equipment. Each diagnostic metric has explicit acceptance criteria to make the test deterministic and repeatable. The 20-minute window accommodates preventive maintenance between shifts without disrupting care schedules. | subsystem, vehicle-platform, sil-1, session-707, idempotency:sub-platform-usbc-diagnostic-707
SUB-VEHICLEPLATFORM-007 | The Chassis Frame SHALL be constructed from aluminium alloy (6061-T6 or equivalent) with a minimum wall thickness of 3mm at structural load points, providing a rated static load capacity of 150kg (user 120kg + equipment 30kg), with a maximum deflection of 2mm under full load, and shall be corrosion-resistant to humidity up to 95% RH for indoor care facility environments. | subsystem-requirements | Vehicle Platform |  | Test | 86 |  | Lint finding: Chassis Frame is classified as Physical Medium but has no material property requirements. 6061-T6 aluminium provides an acceptable strength-to-weight ratio for medical mobility equipment (yield strength 276 MPa). 150kg load capacity with 2:1 safety margin covers the 95th percentile user weight plus equipment. 95% RH corrosion resistance is required for care facility environments including wet areas and cleaning protocols. | subsystem, vehicle-platform, sil-2, session-707, idempotency:sub-platform-chassis-material-707
SUB-VEHICLEPLATFORM-008 | The Electronics Bay SHALL be constructed from steel or aluminium enclosure rated IP54 (dust-protected, splash-resistant), with internal thermal management maintaining component ambient temperature between 0°C and 55°C during operation, and shall withstand vibration levels up to 2g at 5-50Hz without component dislodgement or connector failure. | subsystem-requirements | Vehicle Platform |  | Test | 86 |  | Lint finding: Electronics Bay classified as Physical Medium but has no material property or environmental requirements. IP54 is the minimum for care facility environments where cleaning fluids may be applied nearby. The 0-55°C thermal range matches the operating envelope of COTS embedded processors and power electronics within. 2g vibration at 5-50Hz reflects doorway-crossing and ramp-traversal shock loads in a care facility environment. | subsystem, vehicle-platform, sil-1, session-707, idempotency:sub-platform-ebay-material-707
SUB-VEHICLEPLATFORM-009 | The Chassis Frame SHALL be designed and tested in accordance with ISO 7176-8 (Requirements for static, impact, and fatigue strengths for wheelchairs) for structural adequacy under static loading of 1.5× maximum user mass plus equipment, and shall withstand impact loads of 70J without permanent deformation of safety-critical members. | subsystem-requirements | Vehicle Platform |  | Test | 86 |  | Chassis Frame is classified as Regulated and Structural by UHT Substrate. ISO 7176-8 is the applicable standard for powered wheelchair structural requirements, which is the closest normative framework for a BCI-controlled mobility device. Without a structural compliance requirement, there is no design floor for chassis member sizing against user mass plus overturning loads. 1.5× safety factor and 70J impact threshold are drawn from ISO 7176-8 Table 1 for Class B (indoor/light outdoor) powered wheelchairs, consistent with the care facility operating environment defined in STK-REQ-013. | subsystem, vehicle-platform, compliance, session-708, qc, idempotency:sub-chassis-iso7176-compliance-708
SUB-VEHICLEPLATFORM-010 | After any chassis shock event exceeding 3g vertical acceleration detected by the inclinometer shock channel, the Safety Monitor Processor SHALL assert a structural-integrity-check-required flag, inhibit all vehicle propulsion, and require authorised maintenance acknowledgement before resuming normal operation. | subsystem-requirements | Vehicle Platform |  | Inspection |  |  | A 3g impact at the ISO 7176-8 (Static, impact, and fatigue strengths for wheelchairs) chassis load limit may cause micro-fractures or fastener loosening not visible externally. Inhibiting motion after a 3g shock event prevents operation with potentially compromised structural integrity. The inclinometer shock detection channel on the ST LSM6DSO used in the Perception Subsystem provides the detection mechanism without additional hardware. This defines the safe state for chassis structural limit exceedance identified by the rt-missing-safe-state finding on SUB-REQ-027. | subsystem, structural-subsystem, sil-1, session-717, reqs-eng-session-717, safe-state, addresses:SUB-REQ-027, idempotency:ss-sub-027-chassis-shock-safestate-717
SYS-REQ-001 | The EEG Accessibility Buggy SHALL classify user motor imagery and SSVEP EEG signals with a minimum accuracy of 85% across a 4-class command set (forward, left, right, stop) during Normal Navigation mode, measured over a rolling 2-minute window. | system-requirements | System Requirements |  | Test | 86 |  | Derived from STK-REQ-001 (independent navigation) and STK-REQ-002 (user safety). The 85% threshold is the minimum accuracy at which misclassification rate (<15%) keeps unintended manoeuvres within the obstacle detection recovery envelope. Below 85%, the misclassification rate exceeds the system's ability to correct via obstacle avoidance, creating collision risk (H-002, H-004). | system, bci, performance, session-696, idempotency:sys-bci-accuracy-696
SYS-REQ-002 | When any safety trigger is detected (complete BCI signal loss, imminent collision, system fault, or manual emergency stop button press), the EEG Accessibility Buggy SHALL de-energise drive motors and engage mechanical brakes within 200ms of trigger detection. | system-requirements | System Requirements |  | Test | 86 |  | Derived from STK-REQ-002 (user safety). H-001 (signal loss) and H-002 (collision) both require SIL 2-3 emergency stop. 200ms budget: 50ms detection + 50ms processing + 100ms mechanical actuation. At 6 km/h (1.67 m/s), 200ms yields 0.33m travel — within the 0.5m safety margin to nearest obstacle at trigger point. | system, safety, sil-3, session-696, idempotency:sys-estop-time-696
SYS-REQ-003 | The EEG Accessibility Buggy SHALL limit maximum travel speed to 6 km/h in Normal Navigation mode and 3 km/h in Degraded/Assisted mode. | system-requirements | System Requirements |  | Test | 100 |  | ISO 7176-14 (Power wheelchairs — Maximum speed) permits up to 15 km/h for powered mobility devices on pathways. 6 km/h Normal Navigation cap is a comfortable indoor walking pace, safe for corridor/building use. 3 km/h Degraded/Assisted cap (50% of Normal) provides meaningful mobility while limiting consequence of mis-commands when BCI accuracy is compromised. Aligns with SYS-REQ-022 which specifies the Degraded/Assisted behaviour in detail. | system, performance, session-696, idempotency:sys-speed-limit-696
SYS-REQ-004 | The EEG Accessibility Buggy SHALL provide a minimum of 4 hours continuous operation and 15 km range on a single battery charge at an average speed of 4 km/h with a 120 kg user payload. | system-requirements | System Requirements |  | Test | 100 |  | Derived from STK-REQ-001 (independent navigation) and operational tempo constraint (2-4 sessions/day, 15-45 min each). 4 hours covers a full day's use with margin. 15 km at 4 km/h average matches the Daily Independent Navigation scenario profile. | system, performance, power, session-696, idempotency:sys-range-696
SYS-REQ-005 | When generalised spike-and-wave EEG activity exceeding 3 Hz is detected across all acquisition channels, the EEG Accessibility Buggy SHALL trigger Emergency Stop within 150ms and transmit a BLE emergency alert containing device ID, alert type, and GPS coordinates to the facility emergency system. | system-requirements | System Requirements |  | Test | 86 |  | Derived from STK-REQ-002 (user safety). H-006 (seizure during operation) is rated catastrophic/SIL 3. 150ms budget is tighter than general E-stop (200ms) because seizure-induced EEG patterns can be misclassified as movement commands, causing erratic buggy behaviour before detection completes. The 50ms margin accounts for the unique danger of a non-ambulatory user seizing while in a moving vehicle. | system, safety, sil-3, session-696, idempotency:sys-seizure-detect-696
SYS-REQ-006 | The EEG Accessibility Buggy SHALL detect obstacles and persons within a 2m forward detection zone and 1m lateral detection zone and prevent collision by reducing speed or stopping, with zero false-negative detections for obstacles larger than 50mm height above floor level. | system-requirements | System Requirements |  | Test | 86 |  | Derived from STK-REQ-002 (user safety) and STK-REQ-014 (bystander safety). H-002 (collision) is rated SIL 2. 2m forward range at 6 km/h provides 1.2s stopping margin. 50mm height threshold catches kerbs, feet, and dropped objects while filtering floor texture variations. | system, safety, sil-2, perception, session-696, idempotency:sys-obstacle-detect-696
SYS-REQ-007 | When BCI signal-to-noise ratio drops below the classification threshold for more than 5 seconds, the EEG Accessibility Buggy SHALL transition to Degraded/Assisted mode with speed limited to 3 km/h and BCI simplified to binary go/stop commands. | system-requirements | System Requirements |  | Test | 86 |  | 5-second SNR threshold prevents transient signal glitches from triggering mode transitions, while still catching genuine degradation within a safe reaction time. 3 km/h Degraded speed (50% Normal Navigation cap) aligns with SYS-REQ-003 and SYS-REQ-022. Binary go/stop reduces cognitive load when the BCI classifier cannot reliably discriminate four motor imagery classes. | system, safety, sil-1, bci, session-696, idempotency:sys-degraded-mode-696
SYS-REQ-008 | When rolling 2-minute BCI classification accuracy drops below 70%, the EEG Accessibility Buggy SHALL transition to Degraded/Assisted mode and generate an audible alert to the care attendant. | system-requirements | System Requirements |  | Test | 100 |  | Derived from STK-REQ-003 (fatigue detection). Cognitive Fatigue scenario: accuracy dropped from 85% to 68% over 3 minutes. The 70% threshold with 2-minute rolling window detects this drift pattern with sufficient lead time to prevent dangerous misclassification. H-004 (misclassification) drives this requirement. | system, bci, safety, session-696, idempotency:sys-fatigue-detect-696
SYS-REQ-009 | When the care attendant activates the rear-mounted override switch, the EEG Accessibility Buggy SHALL transfer steering and throttle authority to the joystick within 100ms, ignore BCI movement commands, and maintain obstacle detection in active mode. | system-requirements | System Requirements |  | Test | 100 |  | Derived from STK-REQ-004 (carer override). The 100ms handover time ensures seamless transition during emergency or post-degradation situations. Obstacle detection remains active during override because the carer may be unfamiliar with the environment or distracted by the user's condition. | system, carer, session-696, idempotency:sys-carer-override-696
SYS-REQ-010 | The EEG Accessibility Buggy SHALL have a maximum vehicle footprint of 700mm wide by 1200mm long, an unladen mass not exceeding 80 kg, a user payload capacity of at least 120 kg, and a turning radius not exceeding 1500mm. | system-requirements | System Requirements |  | Inspection | 86 |  | Derived from STK-REQ-013 (facility infrastructure compatibility). 700mm width provides clearance through 900mm DDA-compliant doorways. 1200mm length fits 1100×1400mm lift cabs. 1500mm turning radius enables U-turns in 2000mm wide corridors. 120 kg covers 95th percentile adult wheelchair users. | system, physical, session-696, idempotency:sys-dimensions-696
SYS-REQ-011 | When the onboard inclinometer detects vehicle tilt exceeding 15 degrees in any axis, the EEG Accessibility Buggy SHALL apply brakes and prevent further motion in the direction of the hazardous slope. | system-requirements | System Requirements |  | Test | 100 |  | Derived from STK-REQ-002 (user safety). H-005 (tip-over) is rated SIL 2. 15° tilt threshold provides margin above the 8° maximum ramp gradient (operational constraint) while triggering before the vehicle's centre-of-gravity reaches the tip-over angle (~25° for a 120 kg user at maximum seat height). | system, safety, sil-2, session-696, idempotency:sys-tip-prevention-696
SYS-REQ-012 | When battery cell temperature exceeds 60°C or cell voltage imbalance exceeds 100mV, the EEG Accessibility Buggy SHALL open the battery disconnect relay within 500ms, activate the fire suppression blanket, sound an 85 dB alarm, and transmit a facility emergency alert. | system-requirements | System Requirements |  | Test | 86 |  | Derived from STK-REQ-002 (user safety). H-003 (battery thermal runaway) is rated catastrophic/SIL 3. 60°C cell temperature is the thermal runaway onset threshold for NMC lithium-ion cells per IEC 62133 (Secondary lithium cells — Safety requirements). 500ms response prevents cascade. The non-ambulatory user cannot self-evacuate, making suppression and external alerting critical. | system, safety, sil-3, power, session-696, idempotency:sys-battery-safety-696
SYS-REQ-013 | The EEG Accessibility Buggy SHALL comply with IEC 60601-1-2 (Electromagnetic compatibility — Requirements and tests for medical electrical equipment) for both radiated emissions and immunity, maintaining safe operation in environments containing MRI fringe fields, electrosurgical units, and 50 Hz fluorescent lighting ballast harmonics. | system-requirements | System Requirements |  | Test | 86 |  | Derived from STK-REQ-016 (EMI tolerance). H-007 (EMI-corrupted EEG) is rated SIL 1. Hospital corridors are the primary operating environment and contain aggressive EMI sources. EMC compliance ensures the EEG acquisition chain degrades gracefully rather than producing undetectable classification errors. | system, regulatory, session-696, idempotency:sys-emc-696
SYS-REQ-014 | The EEG Accessibility Buggy SHALL charge from 20% to 100% state of charge within 4 hours when connected to a 230V 50Hz or 120V 60Hz mains supply via the facility charging dock. | system-requirements | System Requirements |  | Test | 100 |  | Derived from STK-REQ-001 (independent navigation) and power constraint. Overnight charging (22:00-06:00) provides 8 hours — 4-hour charge time ensures full charge with margin for late docking. Same-day top-up between sessions is feasible during the 2-hour midday gap. | system, power, session-696, idempotency:sys-charge-time-696
SYS-REQ-015 | The EEG Accessibility Buggy SHALL provide an automated diagnostic test suite accessible via USB-C service port that verifies motor response, brake torque, sensor calibration, battery cell health, and BCI pipeline integrity, completable within 20 minutes. | system-requirements | System Requirements |  | Demonstration | 100 |  | Derived from STK-REQ-006 (diagnostic interface). Weekly Maintenance scenario: technician must inspect 4 buggies in a 2-hour window. 20-minute test suite per buggy allocates the budget correctly. Any buggy failing brake torque or cell imbalance tests is automatically flagged for removal from service. | system, maintenance, session-696, idempotency:sys-diagnostics-696
SYS-REQ-016 | The EEG Accessibility Buggy SHALL support dual-bank flash firmware updates with automatic rollback to the previous verified firmware image if the new image fails post-update integrity verification. | system-requirements | System Requirements |  | Test | 100 |  | Derived from STK-REQ-007 (firmware updates with rollback). IEC 62304 (Medical device software — Software life cycle processes) requires controlled software update mechanisms for Class C software (safety-critical). Dual-bank ensures the device is never bricked by a failed update — critical for fleet availability. | system, software, session-696, idempotency:sys-firmware-696
SYS-REQ-017 | When Emergency Stop or seizure detection triggers, the EEG Accessibility Buggy SHALL transmit a BLE 5.0 iBeacon-compatible alert containing device ID, alert type, and location to the facility BLE gateway within 500ms, and simultaneously notify the user's paired smartphone. | system-requirements | System Requirements |  | Test | 86 |  | Derived from STK-REQ-005 (carer alerts) and Seizure scenario. The non-ambulatory user cannot call for help — the system must autonomously summon assistance. 500ms transmission latency ensures the facility nurse call system receives the alert before the carer would otherwise notice the stationary buggy. | system, communication, safety, session-696, idempotency:sys-ble-alerts-696
SYS-REQ-018 | The EEG Accessibility Buggy SHALL complete power-on self-diagnostics and user-specific EEG calibration within a combined time of 5 minutes, achieving a classification accuracy threshold of at least 80% before entering Normal Navigation mode. | system-requirements | System Requirements |  | Test | 100 |  | Derived from STK-REQ-001 (independent navigation). Startup/Calibration mode: concept defines 30s diagnostics + 3 min calibration = 3.5 min typical, 5 min maximum. 80% calibration threshold is the minimum accuracy at which the 4-class command set is usable — below this, misclassification exceeds 1-in-5 commands. | system, performance, bci, session-696, idempotency:sys-startup-696
SYS-REQ-019 | The EEG Accessibility Buggy SHALL encrypt all EEG biometric data using AES-256 at rest and TLS 1.3 in transit, implement role-based access control for clinical and maintenance interfaces, and maintain an audit log of all data access events. | system-requirements | System Requirements |  | Test | 86 |  | Derived from STK-REQ-018 (GDPR data protection). EEG data is Article 9 special category biometric data under GDPR. AES-256 and TLS 1.3 are the current minimum standards for health data protection. Role-based access prevents maintenance staff accessing clinical data and vice versa. | system, security, regulatory, session-696, idempotency:sys-data-security-696
SYS-REQ-020 | The EEG Accessibility Buggy SHALL use EEG electrode contacts and headset materials certified as biocompatible per ISO 10993-1 (Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process) for prolonged skin contact (greater than 24 hours cumulative daily use), with no materials classified as cytotoxic, sensitising, or irritant in direct contact with the scalp. | system-requirements | System Requirements |  | Inspection | 86 |  | The EEG headset is a skin-contact medical device worn daily for extended sessions by users who may have limited ability to remove it themselves. ISO 10993-1 biocompatibility evaluation is mandatory under IEC 60601-1 for patient-contacting parts. Users with neurological conditions may have reduced skin sensation, increasing risk of contact dermatitis or pressure injury from non-biocompatible materials. This addresses the high-severity lint finding that the system classification includes Biological/Biomimetic traits with no biocompatibility requirements. | system, safety, biocompatibility, sil-2, session-703, idempotency:sys-biocompatibility-703
SYS-REQ-021 | The EEG Accessibility Buggy SHALL be physically housed in a single integrated chassis enclosure providing IP54 environmental protection per IEC 60529 (protection against dust ingress and water splashing from any direction), with all electronics, battery, and BCI processing mounted within a sealed Electronics Bay. | system-requirements | System Requirements |  | Inspection | 86 |  | EEG accessibility buggy is classified without Physical Object trait (hex 50800000) despite being a physical vehicle. This requirement establishes the physical embodiment: a unified housing with IP54 protection for indoor healthcare and light outdoor use. IP54 is the minimum for wheelchair-equivalent environments per IEC 60601-1-11 (home healthcare environments) and ensures electrode/electronics survive routine cleaning. | system, session-706, idempotency:sys-physical-housing-ip54-706
SYS-REQ-022 | While operating in Degraded/Assisted mode due to BCI accuracy below 70%, the EEG Accessibility Buggy SHALL limit maximum speed to 3 km/h (50% of Normal Navigation maximum), maintain full obstacle detection and emergency stop capability, and display a continuous amber alert on the Status LED Array and a text notification on the Display Unit. | system-requirements | System Requirements |  | Test | 86 |  | STK-REQ-004 requires manual override with obstacle detection remaining active but specifies no measurable degraded-mode performance floor. STK-REQ-005 requires distinct alerts on mode transitions. 3 km/h limit (half of full speed) reduces collision energy by 75% and gives the care attendant additional reaction time, which is derived from pedestrian collision safety analysis (EN ISO 13482: safety requirements for personal care robots). Obstacle detection must remain active to prevent hazards during reduced user control. | system, session-706, sil-2, idempotency:sys-degraded-mode-performance-706
SYS-REQ-023 | When the Main Application Processor fails (watchdog timeout exceeding 50ms, voltage rail failure, or CRC fault detected), the Safety Monitor Processor SHALL assert SAFE_STOP, de-energise drive motors, and engage mechanical brakes within 200ms of fault detection, independent of Main Application Processor firmware state. | system-requirements | System Requirements |  | Test | 86 |  | IEC 61508 (Functional safety of E/E/PE safety-related systems) SIL 2 requires independent fault detection and response for System-Essential components. MAP is system-essential (single point of failure for BCI command processing); without a watchdog-monitored independent failover channel, a stuck MAP executing incorrect commands would have no automatic mitigation. The 50ms watchdog timeout and 200ms failsafe response are consistent with SYS-REQ-002 which mandates 200ms total emergency stop latency. This requirement establishes the SYS-level mandate for the independence implemented in ARC-REQ-002 and SUB-REQ-061. | system, safety, sil-2, session-708, qc, idempotency:sys-map-watchdog-failover-708
VER-REQ-040 | Verify SUB-REQ-008: stream 100 pre-annotated EEG epochs (50 containing known EMG bursts, 50 containing known 50Hz mains interference) through the Artifact Rejection Engine. Measure rejection rate on contaminated epochs (≥90% required) and false rejection rate on clean epochs (≤10% required). Process all 100 epochs; confirm end-to-end epoch latency ≤20ms measured from epoch DMA transfer start to clean-epoch availability on output queue. | verification-requirements | Verification Requirements |  | Test | 57 |  | SUB-REQ-008 specifies quantified rejection accuracy and latency. This test uses annotated EEG data with known artefact content to verify rejection performance against both criteria simultaneously. The 50/50 epoch split ensures both false negative and false positive rates are independently verified. | session-704, qc, verification, bci-subsystem, idempotency:ver-artifact-rejection-qc704
VER-REQ-041 | Verify SUB-REQ-009: recruit five representative test subjects (each with a previously completed calibration session). For each subject, run a 10-minute continuous BCI driving session. Extract rolling 2-minute accuracy windows at 30-second intervals. Calculate mean accuracy per subject and across cohort. Pass criterion: all five subjects achieve ≥75% mean classification accuracy across the session; no 2-minute window falls below 60% (single-window floor). | verification-requirements | Verification Requirements |  | Test | 57 |  | SUB-REQ-009 sets the 75% rolling-window classification floor that gates Normal Navigation entry. Multi-subject testing is required because BCI performance varies significantly between users; a single-subject test cannot validate the fleet-wide threshold. The 60% single-window floor protects against brief anomalous windows masking a failing mean. | session-704, qc, verification, bci-subsystem, idempotency:ver-bci-classifier-qc704
VER-REQ-042 | Verify IFC-REQ-011: connect a logic analyser to the inter-process shared memory bus between EEG Acquisition Module and Artifact Rejection Engine. Stream 256 sample/second × 32 channels at rated throughput for 5 minutes. Measure: (a) actual sample delivery rate must be 256 ±2 Hz; (b) maximum inter-process latency measured per epoch must be ≤5ms; (c) replay 1000 epochs and count frames with dropped channels — loss rate must be ≤0.1%. | verification-requirements | Verification Requirements |  | Test | 57 |  | IFC-REQ-011 carries the raw EEG data path from sensor to processing; errors here propagate to all downstream BCI functions. The 0.1% sample loss rate and 5ms latency are the constraints that maintain classification accuracy at the SUB-REQ-009 threshold. Instrumented bus measurement is required because the shared memory interface has no software-visible loss counter. | session-704, qc, verification, bci-subsystem, idempotency:ver-eeg-acq-ifc011-qc704
VER-REQ-103 | Verify SUB-REQ-079: Configure the Main Application Processor to accept BCI-derived navigation commands via external software interface with HMAC-SHA256 authentication. (a) Send 100 valid signed commands and confirm all are accepted and executed. (b) Send 20 commands with tampered signatures and confirm all are rejected and logged with REJECTED_CMD fault code. (c) Send 20 commands with replayed nonces and confirm rejection. Pass: zero valid commands rejected, zero tampered commands accepted, each rejection logged within 50ms. | verification-requirements | Verification Requirements |  | Test | 57 |  | SUB-REQ-079 is the command authentication requirement derived from the cybersecurity mandate (SYS-REQ-019). A hardware-in-the-loop test with deliberate signature tampering and nonce replay is the only method that demonstrates the cryptographic rejection path functions end-to-end without relying on static analysis. Fault log verification confirms non-repudiation. | session-713, validation, bci-subsystem, security, idempotency:ver-sub079-hmac-713
VER-BCIPROCESSINGSUBSYSTEM-001 | The BLE 5.0 EEG headset interface SHALL be verified by: (a) streaming 32-channel 24-bit EEG data at 250Hz for 30 minutes and measuring packet loss rate across 5 runs at 5m and 10m range, confirming < 0.1% at 5m and < 1% at 10m; (b) measuring BLE connection establishment time across 20 trials confirming < 3s; (c) confirming firmware lock on BLE advertisement interval prevents interference with safety BLE channel. | verification-requirements | BCI Processing Subsystem |  | Test | 86 |  | IFC-REQ-001 is the primary BCI data path — packet loss above threshold degrades classification accuracy below safe operational limits. End-to-end streaming test at actual operating ranges is the only method that captures real RF channel conditions (multipath, interference from CMMS Wi-Fi) in a care facility environment. | verification, bci, session-698, idempotency:ver-ble-eeg-headset-698
VER-BCIPROCESSINGSUBSYSTEM-002 | The BCI Processing Subsystem SNR degradation response SHALL be verified by injecting an attenuated noise baseline into the EEG Acquisition Module input, measuring elapsed time from SNR threshold crossing to STOP command on the Drive Subsystem CAN bus. Pass: STOP emitted within 3000ms on ≥10/10 runs; no drive commands emitted during suppression window. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | Direct verification of SIL-3 safe-state transition. CAN bus monitoring provides objective pass/fail evidence. 10-run statistical sample required for SIL-3 safety case. | verification, bci-processing-subsystem, sil-3, session-699, idempotency:ver-sub011-bci-signal-loss-699
VER-BCIPROCESSINGSUBSYSTEM-003 | The BCI Processing Subsystem command pipeline latency SHALL be verified by replaying a known motor imagery EEG epoch at the acquisition input and measuring end-to-end latency from BLE injection to validated CAN command across 50 consecutive epochs. Pass: 95th percentile latency ≤150ms; no individual measurement >200ms. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | Percentile criterion (P95) reflects real-time system requirements — occasional outliers are acceptable but the distribution must satisfy the cognitive feedback loop constraint. 50-epoch sample provides 95% confidence interval for P95 estimate. | verification, bci-processing-subsystem, sil-3, session-699, idempotency:ver-sub010-bci-latency-699
VER-BCIPROCESSINGSUBSYSTEM-004 | The Command Arbitration Module CAN bus interface SHALL be verified by capturing 1000 command frames during a live BCI session. Pass: all frames comply with CAN 2.0B format; status counter increments monotonically; no inter-frame gap >55ms; maximum transmission latency ≤5ms from software send to bus dominant. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | CAN protocol compliance (a,b) and timing (c,d) are testable at integration via bus analyser. Status counter verification directly tests the Drive Subsystem safe-stop detection mechanism. | verification, bci-processing-subsystem, sil-3, session-699, idempotency:ver-ifc015-can-interface-699
VER-BCIPROCESSINGSUBSYSTEM-005 | Verify SUB-REQ-042: Inject out-of-bounds BCI command pattern (velocity > physical maximum) into the BCI Processing Subsystem output and measure time to safe-stop signal assertion. Pass: safe-stop asserted within 100ms. Also block watchdog kick for 250ms and verify safe-stop assertion. | verification-requirements | BCI Processing Subsystem |  | Test | 57 |  | Two-path verification: invalid command detection and watchdog timeout, both required to cover the SUB-REQ-042 safety constraint fully. Hardware fault injection provides realistic test conditions for SIL-3 verification. | verification, bci-processing-subsystem, sil-3, session-702, idempotency:ver-sub042-702
VER-BCIPROCESSINGSUBSYSTEM-006 | Verify the Feature Extraction Processor loads and applies per-user CSP spatial filter matrices from encrypted calibration storage within 5 seconds of user session initialisation. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | SUB-REQ-012 specifies the 5-second calibration load bound; test must confirm timing across five user calibration profiles on production hardware to validate that encrypted storage read latency is within budget at all thermal operating conditions. | session-710, qc, bci-processing, sil-3, idempotency:ver-sub012-calib-load-710
VER-BCIPROCESSINGSUBSYSTEM-007 | Verify the Artifact Rejection Engine watchdog interface causes the BCI Processing Subsystem supervisor to reset the engine and log the event when no valid epoch output is received for more than 500ms. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | SUB-REQ-017 specifies the 500ms watchdog timeout and mandatory supervisor reset; fault injection test must simulate a halted Artifact Rejection Engine, confirm reset is triggered within 600ms, and verify the event record appears in the post-session log with correct timestamp and cause code. | session-710, qc, bci-processing-subsystem, sil-3, idempotency:ver-sub017-watchdog-710
VER-BCIPROCESSINGSUBSYSTEM-008 | Verify the Main Application Processor encrypts EEG biometric data using AES-256 at rest and TLS 1.3 in transit, enforces role-based access control with at least three roles, and maintains an append-only audit log. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | SUB-REQ-039 mandates AES-256/TLS 1.3 encryption and RBAC with three roles; test must confirm stored data is ciphertext via hex dump of encrypted storage, verify TLS 1.3 handshake using a packet capture, attempt access with each role confirming correct privilege boundaries, and verify audit log entries are immutable and include timestamp, identity, and operation type. | session-710, qc, bci-processing-subsystem, security, idempotency:ver-sub039-aes256-security-710
VER-BCIPROCESSINGSUBSYSTEM-009 | Verify the Artifact Rejection Engine operates from the 3.3V logic rail with maximum 250mA peak current draw during active processing and maintains operation for supply voltage variation between 3.0V and 3.6V. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | SUB-REQ-041 specifies the power consumption ceiling and supply tolerance at SIL-3; test must measure peak current draw during maximum processing load using a precision current meter, and sweep supply voltage from 3.0V to 3.6V in 50mV steps confirming no processing error or dropout occurs, as the 3.3V rail is shared with other safety-relevant logic. | session-710, qc, bci-processing-subsystem, power, sil-3, idempotency:ver-sub041-artifact-power-710
VER-BCIPROCESSINGSUBSYSTEM-010 | Verify electrode interface components contacting user scalp meet ISO 10993-1 (Biological evaluation of medical devices, Part 1) biocompatibility requirements: (a) inspect material data sheets confirming each scalp-contact material has been classified per ISO 10993-1 risk framework for prolonged skin contact; (b) measure contact impedance at 1kHz before and after 50 decontamination cycles with specified clinical disinfectant. Pass: all materials carry ISO 10993-1 biocompatibility classification; contact impedance change ≤ 10% after 50 cycles. Fail: any unclassified material; impedance change > 10%. | verification-requirements | BCI Processing Subsystem |  | Test | 57 |  | SUB-REQ-044 mandates ISO 10993-1 (Biological evaluation of medical devices, Part 1) compliance; verification requires material data sheets confirming biocompatibility classification for each scalp-contact material, plus test report showing electrode impedance before and after 50 decontamination cycles with the specified clinical disinfectant, confirming <10% change in contact impedance. | session-710, qc, bci-processing-subsystem, regulatory, biocompatibility, idempotency:ver-sub044-biocompatibility-710
VER-BCIPROCESSINGSUBSYSTEM-011 | Verify the Artifact Rejection Engine occupies no more than 30% of available Feature Extraction Processor CPU cycles during continuous EEG processing at 250 SPS per channel, with deterministic maximum latency of 4ms per processing frame. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | SUB-REQ-066 specifies CPU budget and real-time latency for the SIL-3 artifact rejection module; test must run continuous EEG signal at 250 SPS/channel for 30 minutes, sample CPU utilisation at 10Hz using a hardware performance monitor, confirm 99th-percentile CPU usage ≤30%, and inject frame timing markers to measure maximum frame processing latency confirming ≤4ms worst-case. | session-710, qc, bci-processing-subsystem, sil-3, idempotency:ver-sub066-artifact-cpu-budget-710
VER-BCIPROCESSINGSUBSYSTEM-012 | Verify the Artifact Rejection Engine transfers cleaned 32-channel EEG epochs as float32 arrays in 1-second windows with 50% overlap at a throughput of 2 epochs/second per channel to the Feature Extraction Processor. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | IFC-REQ-012 specifies the BCI pipeline inter-module data contract at SIL-3; test must inject synthetic 32-channel EEG data, capture the output from the Artifact Rejection Engine interface, confirm float32 encoding, 1-second window length, 50% overlap, and sustained 2 epoch/s/channel throughput measured over 100 consecutive epochs with <1% frame drop. | session-710, qc, bci-processing-subsystem, sil-3, idempotency:ver-ifc012-artifact-feat-710
VER-BCIPROCESSINGSUBSYSTEM-013 | Verify the BCI Classifier passes command probability vectors for four navigation classes as float32 with per-class confidence, SNR index, and rolling 2-minute accuracy metric at minimum 2Hz to the Command Arbitration Module. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | IFC-REQ-014 specifies the classifier output format and rate at SIL-3; test must log classifier outputs during a 10-minute BCI session, confirm the vector format (four float32 probabilities, confidence per class, SNR index, accuracy metric), and verify output rate is ≥2Hz at all times including during high-artifact periods. | session-710, qc, bci-processing-subsystem, sil-3, idempotency:ver-ifc014-classifier-arb-710
VER-BCIPROCESSINGSUBSYSTEM-014 | Verify IFC-REQ-013: Configure a test harness injecting synthetic CSP-projected feature vectors at 100Hz from the Feature Extraction Processor to the BCI Classifier. Verify that 1000 consecutive feature vectors are received, decoded, and classification results returned within the 10ms latency budget. Pass: all 1000 vectors received with CRC intact; latency p99 ≤ 10ms measured by timestamp on both ends; no dropped vectors. Fail: any CRC error, dropped vector, or latency exceedance. | verification-requirements | BCI Processing Subsystem |  | Test | 57 |  | IFC-REQ-013 is SIL-3 rated: it carries the BCI feature data that drives all navigation commands. Failure of this interface directly causes loss of BCI control, triggering H-001 (uncontrolled vehicle movement). The 10ms latency budget comes from end-to-end command latency allocation; test harness injection is the only practical method for 1000-iteration statistical validation. | session-711, validation, bci-processing-subsystem, sil-3, idempotency:ver-ifc013-feature-vectors-sil3-711
VER-BCIPROCESSINGSUBSYSTEM-015 | The USB-C Diagnostic Port functionality SHALL be verified by executing the automated diagnostic test suite from a service laptop and measuring: (a) all test categories complete within 20 minutes; (b) pass/fail results individually reported with timestamped log entries; (c) graceful reconnection without data loss after disconnect during a test run. | verification-requirements | BCI Processing Subsystem |  | Test | 71 |  | SUB-REQ-037 is the maintenance technician's primary verification path for fleet-wide safety checks (STK-REQ-006, Weekly Maintenance scenario). If the diagnostic port fails to report complete results within 20 minutes, weekly fleet maintenance cannot be completed within the scheduled window, creating a regulatory compliance gap under EN 12184. | session-713, validation, bci-processing-subsystem, maintenance, idempotency:ver-sub037-usb-diag-713
VER-BCIPROCESSINGSUBSYSTEM-016 | Verify SYS-REQ-007 system-level integration: With the fully integrated EEG Accessibility Buggy operating in Normal Navigation mode at 4 km/h, inject a sustained SNR degradation below the classification threshold via a BCI test signal source. (a) Confirm the system does NOT transition to Degraded/Assisted mode within the first 4.9 seconds of sustained low SNR. (b) Confirm the system transitions to Degraded/Assisted mode within 5.5 seconds of sustained low SNR onset (5-second threshold plus 500ms tolerance). (c) Verify speed is clamped to 3 km/h within 200ms of mode transition. (d) Verify BCI command set reduces to binary go/stop — inject a left-turn motor imagery pattern and confirm it is NOT executed. (e) Restore SNR above threshold and verify the system returns to Normal Navigation mode within 10 seconds. Fail: transition before 4.9s, transition after 5.5s, speed exceeds 3 km/h, or non-binary command accepted. | verification-requirements | BCI Processing Subsystem |  | Test | 57 |  | Validation session 713 identified that SYS-REQ-007 had no direct system-level VER entry — verification was only through subsystem-level tests (SUB-REQ-011 via VER-REQ-014). This integration test confirms the end-to-end chain: BCI SNR detection → mode transition → speed clamping → command simplification operates correctly when all subsystems are integrated. The 5-second timing boundary is the critical safety parameter — if the system transitions too early, it causes unnecessary mode changes; too late, it allows degraded BCI commands at full speed. | session-714, review, integration-test, bci-processing
VER-BCIPROCESSINGSUBSYSTEM-017 | Verify SYS-REQ-008 system-level integration: With the fully integrated EEG Accessibility Buggy operating in Normal Navigation mode, inject a declining BCI accuracy profile via a test signal source so that rolling 2-minute accuracy crosses the 70% threshold. (a) Confirm the system does NOT transition to Degraded/Assisted mode while rolling accuracy remains at 71% or above. (b) Confirm the system transitions to Degraded/Assisted mode within 2 seconds of rolling accuracy dropping to 69% or below. (c) Verify an audible alert of minimum 65 dBA is generated within 1 second of mode transition. (d) Verify speed is limited to 3 km/h. (e) Restore accuracy above 70% for a full 2-minute window and verify Normal Navigation mode resumes. Fail: premature transition, transition delayed beyond 2 seconds, no audible alert, speed exceeds 3 km/h in Degraded mode, or no recovery after accuracy restoration. | verification-requirements | BCI Processing Subsystem |  | Test | 57 |  | Validation session 713 identified that SYS-REQ-008 had no direct system-level VER entry — verification was only at the subsystem level (SUB-REQ-009 via VER-REQ-041). This integration test confirms the end-to-end cognitive fatigue detection chain: rolling accuracy calculation → threshold crossing → mode transition → audible alert → speed limitation. The 70% boundary is the critical decision point — below it, the BCI classifier is unreliable and the vehicle must restrict autonomy. | session-714, review, integration-test, bci-processing
VER-COMMUNICATIONSUBSYSTEM-001 | The Communication Subsystem cellular modem network isolation SHALL be verified by injecting UDP packets addressed to the CAN bus gateway IP range from the modem interface while monitoring all internal bus traffic. Pass: zero packets from the cellular modem interface appear on CAN bus segments across 100 injection attempts. | verification-requirements | Communication Subsystem |  | Test | 71 |  | Active injection test from the modem interface is the only reliable method to confirm the firewall blocks CAN-addressed traffic at runtime. Static configuration inspection alone is insufficient for a security-critical requirement. | verification, communication-subsystem, session-701, idempotency:ver-sub031-701
VER-COMMUNICATIONSUBSYSTEM-002 | Verify SUB-REQ-038: Trigger simulated Emergency Stop event in the Communications Controller and measure BLE advertisement receipt timestamp at BLE gateway sniffer. Pass: alert received within 500ms of trigger for 10 consecutive tests with gateway at 10m range. | verification-requirements | Communication Subsystem |  | Test | 57 |  | End-to-end BLE timing test from trigger to gateway receipt is the only complete verification of the 500ms alert latency chain, accounting for BLE advertisement interval jitter and radio propagation delay. | verification, communication-subsystem, session-702, idempotency:ver-sub038-702
VER-COMMUNICATIONSUBSYSTEM-003 | Verify IFC-REQ-027: connect Communication Controller and Cellular Modem via USB, issue AT+CGDCONT APN configuration command, initiate LTE data session, and measure round-trip ping latency to a test server over LTE. Pass: USB enumeration succeeds as CDC-ECM, APN configured, ping RTT <= 100ms in 95th percentile over 100 measurements. | verification-requirements | Communication Subsystem |  | Test | 57 |  | Integration test confirming the CDC-ECM USB interface is correctly enumerated and the latency budget supports the telemetry uplink requirement SUB-REQ-046. | verification, communication-subsystem, session-703, idempotency:ver-ifc027-703
VER-COMMUNICATIONSUBSYSTEM-004 | Verify IFC-REQ-028: pair BLE module with EEG headset reference device, stream 8-channel EEG data for 60 seconds, timestamp each packet at BLE receive and at BCI Acquisition Module input. Pass: connection interval 7.5ms confirmed in BLE sniffer trace, packet loss < 0.1%, end-to-end latency <= 20ms in 99th percentile. | verification-requirements | Communication Subsystem |  | Test | 57 |  | Confirms BLE 5.2 connection parameters and validates the 20ms wireless transport budget allocated from the 50ms BCI end-to-end latency SYS-REQ-001. | verification, communication-subsystem, session-703, idempotency:ver-ifc028-703
VER-COMMUNICATIONSUBSYSTEM-005 | Verify SUB-REQ-067: Trigger a simulated Seizure Emergency Stop on the EEG Accessibility Buggy while the facility emergency system BLE receiver is active. Measure time from E-Stop trigger to BLE alert packet receipt at the receiver. Pass if packet received within 500ms with correct device ID, alert type code, GPS coordinates, and UTC timestamp on 10 consecutive trials. | verification-requirements | Communication Subsystem |  | Test | 57 |  | Integration test at Communication Subsystem boundary to verify BLE alert timing and content on emergency trigger. 10 trials provide statistical confidence at low failure rate. GPS coordinate validation confirms the facility can dispatch to correct location. | verification, communication-subsystem, session-707, idempotency:ver-sub-req-067-707
VER-COMMUNICATIONSUBSYSTEM-006 | Verify the Bluetooth LE Module maintains a stable BLE 5.2 connection sustaining 32-channel EEG data at 8kHz sample rate (512kbps) within a 3m operating range with no physical obstructions. | verification-requirements | Communication Subsystem |  | Test | 71 |  | SUB-REQ-030 specifies the BLE throughput and range requirement; test must stream 32-channel EEG data at 8kHz from a calibrated EEG headset simulator at 3m range, log packet loss and connection drops over 60 minutes, and confirm throughput remains ≥512kbps with <0.1% packet loss. | session-710, qc, communication-subsystem, sil-1, idempotency:ver-sub030-ble-throughput-710
VER-COMMUNICATIONSUBSYSTEM-007 | Verify the Cellular Modem authenticates to the remote telemetry server using mutual TLS 1.3 with a device-unique X.509 certificate, rejecting server certificates not signed by the system root CA. | verification-requirements | Communication Subsystem |  | Test | 71 |  | SUB-REQ-045 mandates mTLS with device-unique X.509 provisioned at manufacture; test must attempt connection with a valid server certificate (confirm acceptance), a self-signed certificate (confirm rejection), and an expired certificate (confirm rejection), logging TLS handshake state for each attempt. | session-710, qc, communication-subsystem, idempotency:ver-sub045-mtls-710
VER-COMMUNICATIONSUBSYSTEM-008 | Verify the Cellular Modem transmits session telemetry to the remote server at a minimum interval of 10 seconds with a maximum uplink latency of 2 seconds under nominal 4G coverage. | verification-requirements | Communication Subsystem |  | Test | 71 |  | SUB-REQ-046 specifies 10-second transmission interval and 2-second uplink latency; test must capture 100 consecutive telemetry frames at a nominal 4G test network, measure inter-frame interval (confirm all ≤10s), and measure round-trip server acknowledgment time (confirm all ≤2s at 95th percentile). | session-710, qc, communication-subsystem, idempotency:ver-sub046-telemetry-rate-710
VER-COMMUNICATIONSUBSYSTEM-009 | Verify IFC-REQ-005: Connect a service laptop via USB-C cable and execute firmware flash, configuration export, fault log download, and calibration data upload operations. Pass: (a) each operation completes successfully without data corruption (hash check); (b) firmware flash completes within 10 minutes; (c) fault log export produces structured JSON with all required fields; (d) calibration data import round-trips without any field alteration. Fail: any operation fails, hash mismatch, or log missing required fields. | verification-requirements | Communication Subsystem |  | Test | 57 |  | IFC-REQ-005 is the primary maintenance interface; failure to reliably perform any of the four operations would block the Weekly Maintenance scenario and prevent the facility technician from executing firmware updates (firmware rollback safety) or reviewing fault history for clinical compliance. | session-711, validation, communication-subsystem, maintenance, idempotency:ver-ifc005-usbc-service-711
VER-COMMUNICATIONSUBSYSTEM-010 | Verify SUB-REQ-056: Connect a test host to the Communication Subsystem USB-C service port. Issue each motor response verification command individually from the defined diagnostic API. Measure time from command issuance to pass/fail result with logged timestamp. Pass: all commands return a result with timestamp within 10 seconds of issuance on each of 10 consecutive command cycles. Fail: any result arrives after 10 seconds or lacks a timestamp. | verification-requirements | Communication Subsystem |  | Test | 57 |  | SUB-REQ-056 is the communication subsystem's interface to the maintenance diagnostic chain (SYS-REQ-015, STK-REQ-006). The 10-second response window is the per-test budget within the 20-minute fleet maintenance schedule. Testing all 10 cycles confirms reliability across the complete inspection round, not just first-call behaviour. | session-713, validation, communication-subsystem, maintenance, idempotency:ver-sub056-usbc-api-713
VER-COMMUNICATIONSUBSYSTEM-011 | Verify IFC-REQ-004: Pair the vehicle with a test smartphone running the companion app via BLE 5.0. (a) Confirm status display (operating mode, battery level, session history) updates at ≥1 Hz. (b) Trigger a simulated emergency stop and verify BLE emergency alert reaches the companion app within 500ms, measured from E-stop relay actuation to app notification. (c) Capture BLE traffic with a sniffer and confirm all advertising and connection PDUs use LE Secure Connections pairing with AES-128 link-layer encryption (no legacy pairing). Pass: all three criteria met. Fail: update rate <1Hz, alert latency >500ms, or unencrypted PDUs detected. | verification-requirements | Communication Subsystem |  | Test | 57 |  | IFC-REQ-004 is the companion app interface, critical for carer and user situational awareness outside direct line-of-sight. Encryption verification by traffic capture is the only reliable method to confirm BLE security without trusting firmware attestation alone — MHRA Class IIb audit may require this evidence. | session-713, validation, communication-subsystem, idempotency:ver-ifc004-ble-companion-713
VER-COMMUNICATIONSUBSYSTEM-012 | Verify IFC-REQ-006: Configure a test CMMS server endpoint on the facility Wi-Fi (802.11ac) network with TLS 1.3 and API key authentication. (a) Trigger a maintenance event and confirm a REST API POST is received by the test server within 30 seconds with the correct JSON schema (maintenance log, fault code, fleet ID). (b) Configure an invalid API key on the server and confirm the vehicle logs an CMMS_AUTH_FAIL event and ceases retry attempts after 3 failures. (c) Capture Wi-Fi traffic and confirm all HTTP traffic is TLS 1.3 — no cleartext requests. Pass: all three criteria met. | verification-requirements | Communication Subsystem |  | Test | 57 |  | IFC-REQ-006 connects the vehicle to the facility CMMS, which manages regulatory maintenance records (STK-REQ-012). Authentication failure handling prevents credential stuffing. TLS 1.3 verification by traffic capture provides non-repudiable evidence for MHRA post-market surveillance compliance. | session-713, validation, communication-subsystem, maintenance, idempotency:ver-ifc006-cmms-rest-713
VER-COMMUNICATIONSUBSYSTEM-013 | Verify IFC-REQ-025: Connect a UART analyser to the Communication Controller to Bluetooth LE Module UART link. (a) Issue a sequence of 500 HCI commands at maximum throughput and measure end-to-end transfer latency — Pass: ≤5ms latency for 95th percentile across 500 commands. (b) Disable hardware flow control (RTS/CTS) lines and confirm data corruption is detected — Pass: module reports HCI error code within 100ms of buffer overrun. (c) Measure actual baud rate using frequency counter — Pass: 1.000 Mbps ±0.5%. Fail: latency >5ms at P95, error not detected within 100ms, or baud error >0.5%. | verification-requirements | Communication Subsystem |  | Test | 57 |  | IFC-REQ-025 carries BLE HCI traffic for all wireless alerts including seizure emergency notifications. Latency >5ms delays BLE advertising, potentially extending the gap between emergency stop and facility alert. Hardware flow control test confirms UART configuration is not silently operating without it, which would cause undetected packet loss in the alert path. | session-713, validation, communication-subsystem, idempotency:ver-ifc025-uart-ble-713
VER-DRIVESUBSYSTEM-001 | Verify IFC-REQ-018: Apply 20kHz PWM gate-drive signals to the Left Drive Motor Assembly and measure encoder feedback period at 10kHz sample rate with an oscilloscope. Pass criteria: PWM frequency within ±1% of 20kHz, encoder signal received within 1ms propagation delay, and velocity closed-loop error converges to within ±0.1 km/h in steady state. | verification-requirements | Drive Subsystem |  | Test | 57 |  | Integration test to verify interface compliance at the motor-controller boundary. Encoder propagation delay and PWM accuracy are testable on a hardware-in-the-loop bench before vehicle assembly. | verification, drive-subsystem, session-700, idempotency:ver-ifc018-700
VER-DRIVESUBSYSTEM-002 | Verify IFC-REQ-019: Apply 20kHz PWM gate-drive signals to the Right Drive Motor Assembly and measure encoder feedback period at 10kHz sample rate. Pass criteria: symmetric with IFC-REQ-018 verification — PWM frequency within ±1% of 20kHz, encoder delay within 1ms, velocity error within ±0.1 km/h. Additionally verify left-right encoder latency difference does not exceed 0.5ms. | verification-requirements | Drive Subsystem |  | Test | 57 |  | Differential drive correctness requires matched left-right encoder timing. A >0.5ms asymmetry translates to >3cm/s differential at max speed, causing unacceptable yaw error during straight-line travel. | verification, drive-subsystem, session-700, idempotency:ver-ifc019-700
VER-DRIVESUBSYSTEM-003 | Verify IFC-REQ-020: Inject an overcurrent condition on the Drive Power Stage test bench and measure the time from overcurrent event to logic-level trip signal at the Motor Controller Unit. Pass criteria: trip signal received within 5ms, pre-charge sequence completes without bus voltage overshoot exceeding 5% above 48V nominal, MCU logs a fault code within 10ms of receiving trip signal. | verification-requirements | Drive Subsystem |  | Test | 57 |  | Hardware overcurrent trip latency is a safety-critical timing requirement. Test bench injection ensures measurement independence from operating load. Fault logging within 10ms ensures diagnostic traceability for maintenance and incident review. | verification, drive-subsystem, session-700, idempotency:ver-ifc020-700
VER-DRIVESUBSYSTEM-004 | The Motor Controller Unit CAN heartbeat watchdog SHALL be verified by injecting a CAN heartbeat loss and measuring elapsed time to zero motor velocity. Pass: both motors reach zero velocity within 250ms, MCU remains in safe state until a valid restart CAN frame is received, and no uncommanded restart occurs within a 30-second observation window. | verification-requirements | Drive Subsystem |  | Test | 71 |  | Safe-state transition is a SIL 2 safety requirement. End-to-end timing must be measured on the actual vehicle hardware, not simulated, to credit it in the safety case. 30-second observation window for spontaneous restart is the minimum dwell time to detect latent firmware restart bugs. | verification, drive-subsystem, session-700, sil-2, idempotency:ver-sub020-safe-state-700
VER-DRIVESUBSYSTEM-005 | Verify SUB-REQ-058: Suppress MCU CAN heartbeat and measure time to coast-to-stop state and MCU_FAULT assertion to Safety Monitor Processor. Pass: coast-to-stop commanded within 100 ms, MCU_FAULT asserted within 50 ms of timeout detection, 10 trials. | verification-requirements | Drive Subsystem |  | Test | 57 |  | SUB-REQ-058 is a SIL-2 redundancy requirement; injected CAN failure in a hardware-in-the-loop setup verifies the timing without degrading the production system. | verification, drive-subsystem, session-705, sil-2, idempotency:ver-sub058-705
VER-DRIVESUBSYSTEM-006 | The EEG Accessibility Buggy SHALL be verified to confirm that in Care Attendant Override mode the Drive Subsystem responds to joystick inputs within 100ms, obstacle detection remains active with ≤200ms response, maximum speed is limited to 6 km/h, and the system sustains override operation for ≥30 minutes at 50% throttle. | verification-requirements | Drive Subsystem |  | Test | 86 |  | Provides structured test procedure for SUB-REQ-076 override mode performance. Test covers four distinct measurable criteria: latency, obstacle detection continuity, speed ceiling, and endurance. Conducted using calibrated timing equipment in a controlled corridor environment simulating care facility use. | verification, drive, session-708, qc, idempotency:ver-drive-override-mode-perf-708
VER-DRIVESUBSYSTEM-007 | Verify the Motor Controller Unit implements closed-loop velocity control at 100 Hz with 1024 PPR quadrature encoder feedback, achieving steady-state velocity error not exceeding ±0.1 km/h at any commanded speed. | verification-requirements | Drive Subsystem |  | Test | 71 |  | SUB-REQ-018 specifies 100 Hz update rate and ±0.1 km/h steady-state error; test must command five speed set-points across the operating range, measure steady-state error with a calibrated tachometer, and log PID update rate using a hardware timer capture to confirm 100 Hz scheduling compliance. | session-710, qc, drive-subsystem, sil-2, idempotency:ver-sub018-velocity-control-710
VER-DRIVESUBSYSTEM-008 | Verify the Motor Controller Unit enforces the 6 km/h Normal Navigation mode speed limit and 2 km/h Restricted mode speed limit by clamping velocity commands exceeding each threshold before motor drive stage application. | verification-requirements | Drive Subsystem |  | Test | 71 |  | SUB-REQ-019 specifies hard speed limits in two operating modes; test must inject velocity commands 20% above each limit, measure vehicle speed with a calibrated roller dynamometer, and confirm the clamped output does not exceed the specified limit, independent of BCI or joystick input value. | session-710, qc, drive-subsystem, sil-2, idempotency:ver-sub019-speed-limit-710
VER-DRIVESUBSYSTEM-009 | Verify the Drive Power Stage trips the hardware overcurrent protection and isolates motor phase outputs within 5ms when per-channel motor current exceeds 30A, independent of Motor Controller Unit firmware. | verification-requirements | Drive Subsystem |  | Test | 71 |  | SUB-REQ-021 is a safety-critical hardware protection requirement at SIL-2; test must inject a controlled overcurrent fault via a current-limited power supply and measure phase output isolation time with an oscilloscope (<5ms), then repeat the test with the MCU firmware disabled to confirm firmware independence. | session-710, qc, drive-subsystem, sil-2, safety, idempotency:ver-sub021-overcurrent-710
VER-DRIVESUBSYSTEM-010 | Verify the Left and Right Drive Motor Assemblies each deliver a minimum continuous shaft output of 250W at 48V DC across 0°C to 40°C with traction maintained on level indoor surfaces up to 5% gradient. | verification-requirements | Drive Subsystem |  | Test | 71 |  | SUB-REQ-022 specifies per-motor 250W continuous output and gradient traction; each motor must be tested individually on a dynamometer at 0°C, 25°C, and 40°C, recording shaft power and confirming 250W minimum; traction test on 5% inclined surface confirms no wheel slip at full load. | session-710, qc, drive-subsystem, idempotency:ver-sub022-motor-power-710
VER-DRIVESUBSYSTEM-011 | Verify the Motor Controller Unit hardware overcurrent protection disconnects motor phase outputs within 5ms when phase current exceeds 120% of rated maximum, with no firmware execution required. | verification-requirements | Drive Subsystem |  | Test | 71 |  | SUB-REQ-064 is a SIL-2 hardware safety function requiring firmware-independent operation; fault injection test must apply a controlled phase current equal to 121% of rated maximum with MCU firmware halted, and measure time from current threshold to phase output de-energisation using an oscilloscope, confirming ≤5ms disconnect in all three test runs. | session-710, qc, drive-subsystem, sil-2, idempotency:ver-sub064-mcu-hw-overcurrent-710
VER-DRIVESUBSYSTEM-012 | Verify the Drive Subsystem enforces a maximum speed cap of 3 km/h in Degraded/Assisted mode by limiting PWM duty cycle to 50% of Normal Navigation maximum, independent of BCI or joystick input value. | verification-requirements | Drive Subsystem |  | Test | 71 |  | SUB-REQ-065 specifies the SIL-2 degraded mode speed cap applied firmware-side before MCU; test must command full joystick and full BCI input while in Degraded mode and measure vehicle speed on a roller dynamometer, confirming it does not exceed 3 km/h; additionally verify via oscilloscope that PWM duty cycle does not exceed 50% of the Normal Navigation maximum. | session-710, qc, drive-subsystem, sil-2, idempotency:ver-sub065-degraded-speed-cap-710
VER-DRIVESUBSYSTEM-013 | Verify SUB-REQ-077: Submit Drive Subsystem assembly (Motor Controller Unit and Motor Power Isolation Relay) for IEC 60601-1-2 (Medical Electrical Equipment — Electromagnetic disturbances) conducted and radiated emissions and immunity testing by an accredited test laboratory. Pass: test report from accredited lab confirms compliance with all applicable limits and no performance degradation during immunity tests. Fail: any emissions exceed limits or immunity test causes motor controller malfunction. | verification-requirements | Drive Subsystem |  | Test | 57 |  | SUB-REQ-077 mandates IEC 60601-1 (Medical electrical equipment, Part 1: General requirements for basic safety) compliance for the drive subsystem; the 48V BLDC motor controller is a significant EMC source that could cause H-007 interference with the EEG acquisition subsystem if emissions exceed the medical device standard limits. | session-711, validation, drive-subsystem, compliance, idempotency:ver-sub077-iec60601-drive-711
VER-HMISUBSYSTEM-001 | Verify SUB-REQ-032: trigger E-stop via software command and measure Audio Alert Module SPL with a calibrated sound level meter at 1m. Record onset time from E-stop signal assertion to first audible tone. Pass: SPL >= 80dB at 1m, onset <= 100ms, sustained tone until E-stop cleared, in 5 consecutive tests. | verification-requirements | HMI Subsystem |  | Test | 57 |  | Sound level meter measurement is the standard method for verifying audible alert compliance per IEC 60601-1-8 (Medical electrical equipment — Alarm systems). | verification, hmi-subsystem, session-701, idempotency:ver-sub032-701
VER-HMISUBSYSTEM-002 | The Carer Override Module switch activation latency SHALL be verified using a hardware override switch simulator, measuring the timestamp delta from switch assertion to first joystick command accepted by the Drive Subsystem via CAN bus analyser. Pass: delta ≤100ms on 10 consecutive trials. | verification-requirements | HMI Subsystem |  | Test | 71 |  | Hardware-in-the-loop timing test with CAN bus analysis provides unambiguous measurement of authority transfer latency across the HMI/Drive interface. | verification, hmi-subsystem, session-702, idempotency:ver-sub034-702
VER-HMISUBSYSTEM-003 | Verify IFC-REQ-029: inject alert-code 0x01 (Emergency Stop) via I2C at 400kHz from the application processor, measure time from I2C transaction completion to first audio sample output from Audio Alert Module. Pass: onset latency <= 50ms in 10 consecutive trials, tone frequency and duration match alert-code-0x01 specification. | verification-requirements | HMI Subsystem |  | Test | 57 |  | Validates the I2C interface latency and Audio Alert Module tone-pattern firmware against the 50ms onset budget derived from SYS-REQ-002 Emergency Stop response chain. | verification, hmi-subsystem, session-703, idempotency:ver-ifc029-703
VER-HMISUBSYSTEM-004 | Verify SUB-REQ-047: power up the display unit, enter Normal Navigation mode with a test BCI signal source, and measure display update frequency for BCI accuracy, speed, battery SoC, and mode fields over 30 seconds. Measure screen brightness with a calibrated photometer at 0.5m distance. Pass: all four fields update >= 2Hz, brightness >= 300 cd/m2. | verification-requirements | HMI Subsystem |  | Test | 57 |  | Validates display content completeness and refresh rate required for real-time care attendant monitoring per SYS-REQ-008 and SYS-REQ-018 readiness criteria. | verification, hmi-subsystem, session-703, idempotency:ver-sub047-703
VER-HMISUBSYSTEM-005 | Verify SUB-REQ-048: mount Status LED Array on vehicle rear at operational position. Using a calibrated photometer, measure luminous intensity at 11 angular positions spanning 120 degrees (0°, ±12°, ±24°, ±36°, ±48°, ±60°) at a distance of 1m in a darkened environment. Pass criterion: all 11 measurements ≥10 mcd per LED element for each of the three colour states (green, amber, red). Repeat at low battery (battery at 10% SoC) to confirm intensity is maintained. | verification-requirements | HMI Subsystem |  | Test | 57 |  | SUB-REQ-048 specifies a 120-degree arc visibility and 10 mcd minimum intensity to ensure care attendants can read vehicle state from any normal approach angle. Grid measurement at 12-degree intervals ensures no null points exist within the specified arc. Low-battery verification ensures illumination is maintained when it is most needed (approaching shutdown). | session-704, qc, verification, hmi-subsystem, idempotency:ver-led-visibility-qc704
VER-HMISUBSYSTEM-006 | Verify IFC-REQ-030: using SPI analyser on the MAP–LED Array bus, command a mode transition (Normal→Degraded) and measure elapsed time from transition command issue to first valid LED frame on SPI bus. Pass criterion: ≤20ms latency on 20 consecutive transitions. Confirm the frame carries independently addressable RGB values for minimum 6 elements by inspecting SPI frame structure against protocol specification. | verification-requirements | HMI Subsystem |  | Test | 57 |  | IFC-REQ-030 sets a 20ms LED frame latency to ensure state changes are communicated to care attendants within one control cycle. The SPI protocol also carries per-element RGB addressing which is required for status encoding. Frame-level inspection verifies the protocol is implemented as specified and not simplified to a single colour bus. | session-704, qc, verification, hmi-subsystem, idempotency:ver-led-spi-ifc030-qc704
VER-HMISUBSYSTEM-007 | The Carer Override Module rear switch activation latency SHALL be verified at 4 km/h vehicle speed, measuring elapsed time from switch edge to first accepted joystick command via oscilloscope on the CARER_OVERRIDE line and Drive Subsystem CAN bus. Pass: elapsed time ≤100ms across 20 consecutive activations; no single measurement >120ms. | verification-requirements | HMI Subsystem |  | Test | 71 |  | Oscilloscope measurement at physical signal boundaries gives unambiguous timing evidence independent of software logging latency. | verification, hmi-subsystem, session-706, idempotency:ver-sub062-706
VER-HMISUBSYSTEM-008 | Verify SUB-REQ-068: Inject a BCI accuracy threshold crossing event (accuracy dropping from 75% to 65%) while measuring Audio Alert Module response. Pass if 880Hz tone at >=85 dBSPL at 1m is audible within 200ms of threshold crossing and sustained until mode transition or HMI acknowledgement, on 5 consecutive trials. | verification-requirements | HMI Subsystem |  | Test | 57 |  | Functional test of HMI Audio Alert Module response time and sound level. Using a sound level meter at 1m confirms the 85 dBSPL specification is met in situ. 5 trials give confidence without excessive repetition for an audio test. | verification, hmi-subsystem, session-707, idempotency:ver-sub-req-068-707
VER-HMISUBSYSTEM-009 | Verify SUB-REQ-069: Activate the rear-mounted override switch while the buggy is in BCI navigation mode. Measure time from switch activation to joystick steering authority confirmed active and BCI commands confirmed disabled. Pass if handover confirmed within 100ms and BCI commands produce no motion response, on 10 consecutive trials including 3 during active BCI movement command. | verification-requirements | HMI Subsystem |  | Test | 57 |  | Tests the joystick override handover timing at the HMI Subsystem boundary. 10 trials including 3 during active BCI command stress-tests the interlock logic to confirm authority transfer is reliable even when motion commands are in flight. Pass criteria directly verify SYS-REQ-009 derivation. | verification, hmi-subsystem, session-707, idempotency:ver-sub-req-069-707
VER-HMISUBSYSTEM-010 | Verify SUB-REQ-073: Inject a BCI accuracy drop to 65% (degraded mode trigger). Observe Status LED Array and Display Unit. Pass if amber 2Hz pulse is visible within 100ms of mode entry and Display Unit shows the specified text string within 100ms, both persisting until mode exit, on 5 consecutive trials. | verification-requirements | HMI Subsystem |  | Demonstration | 57 |  | Demonstration test for visual HMI feedback in degraded mode. 5 trials confirm consistency of the 100ms response latency. Text string exact match ensures care staff receive unambiguous mode notification. | verification, hmi-subsystem, session-707, idempotency:ver-sub-req-073-707
VER-HMISUBSYSTEM-011 | Verify the Status LED Array displays the correct distinct colour state for each operating mode and completes state transitions within 200ms of the triggering event. | verification-requirements | HMI Subsystem |  | Demonstration | 71 |  | SUB-REQ-033 specifies five distinct LED states (Normal/Degraded/E-stop/Charging/Calibration) and 200ms transition time; verification must step through all five modes in sequence, photographically confirm the correct colour state is active, and measure transition time with an oscilloscope trigger from the triggering event to LED change. | session-710, qc, hmi-subsystem, idempotency:ver-sub033-led-states-710
VER-HMISUBSYSTEM-012 | Verify SUB-REQ-052: Activate the rear-mounted override switch and measure elapsed time from switch edge to CARER_OVERRIDE signal assertion at the Safety Monitor Processor input pin, using an oscilloscope trigger on the switch and SMP GPIO. Pass: assertion time ≤ 50ms on all 20 consecutive activations; signal maintained continuously while switch is held. Fail: any measurement > 50ms or signal drop during hold. | verification-requirements | HMI Subsystem |  | Test | 57 |  | SUB-REQ-052 is in the carer override safety chain; delayed assertion of CARER_OVERRIDE could allow unintended BCI-driven movement while the carer attempts to take control, as seen in the Signal Degradation scenario where carer intercepts the user in the EMI zone. | session-711, validation, hmi-subsystem, idempotency:ver-sub052-carer-override-assert-711
VER-HMISUBSYSTEM-013 | The Carer Override Module joystick command latency SHALL be verified with CARER_OVERRIDE asserted by measuring elapsed time from joystick deflection input to CAN frame transmission on the Drive Subsystem bus via logic analyser. Pass: CAN frame transmitted within 20ms on ≥19/20 activations; command rate ≥50Hz sustained over 10 seconds. | verification-requirements | HMI Subsystem |  | Test | 71 |  | SUB-REQ-053 specifies the carer joystick command routing latency during override; 20ms latency maps to 50Hz command rate needed for smooth carer-controlled navigation in the Signal Degradation scenario where the carer guides through an EMI zone. | session-711, validation, hmi-subsystem, idempotency:ver-sub053-joystick-routing-711
VER-HMISUBSYSTEM-014 | Verify IFC-REQ-026: Connect an SPI logic analyser to the Main Application Processor to Display Unit link. (a) Command display to render a full-frame 800x480 16-bit image and measure SPI clock frequency — Pass: 40 MHz ±2%. (b) Render 60 consecutive frames and measure frame rate — Pass: ≥30 fps with no frame drops (chip-select transitions counted). (c) Send a chip-select pulse while bus is idle and confirm display responds without bus contention. Fail: clock frequency out of tolerance, frame rate <30fps, or bus contention detected. | verification-requirements | HMI Subsystem |  | Test | 57 |  | IFC-REQ-026 drives the HMI display that shows operating mode, battery level, and alerts to care attendants. Below 30fps the display appears to freeze during navigation, delaying carer recognition of mode transitions. SPI frequency verification confirms the processor has not misconfigured the clock divider post-firmware-update. | session-713, validation, hmi-subsystem, idempotency:ver-ifc026-spi-display-713
VER-PERCEPTIONSUBSYSTEM-001 | Verify SUB-REQ-023: place calibrated targets at 0.1m, 0.5m, 1.0m, 1.5m, and 2.0m in the forward arc under lighting conditions from 0 lux (IR illumination only) to 10,000 lux. Confirm distance readings are within 50mm of ground truth at 10Hz sustained over 30 seconds. Pass: all five distances within tolerance at all three lighting levels. | verification-requirements | Perception Subsystem |  | Test | 57 |  | Physical bench test with calibrated distance targets is the only method that directly validates both the sensor accuracy and the update rate under the specified environmental range. | verification, perception-subsystem, session-701, idempotency:ver-sub023-701
VER-PERCEPTIONSUBSYSTEM-002 | Verify IFC-REQ-023: inject an obstacle scenario via test fixture and capture the SPI bus between the Perception MCU and Safety Monitor Processor with a logic analyser. Confirm 4-byte frame format, 10Hz nominal rate, rolling counter incrementing correctly, and that after two missed frames the Safety Monitor Processor asserts communication fault within 250ms. Pass: all frame format, rate, counter, and fault assertions met in 100 consecutive frames. | verification-requirements | Perception Subsystem |  | Test | 57 |  | Logic analyser capture provides ground-truth verification of the SPI protocol against the IFC-REQ-023 specification without relying on software-side assertions. | verification, perception-subsystem, session-701, idempotency:ver-ifc023-701
VER-PERCEPTIONSUBSYSTEM-003 | Verify SUB-REQ-026: with system running, halt the Perception MCU firmware (power off the MCU or inject a software lockup). Monitor Safety Monitor Processor output. Confirm E-stop sequence is initiated within 150ms of last valid frame. Pass: E-stop asserted within 150ms in 5 consecutive tests from normal operating state. | verification-requirements | Perception Subsystem |  | Test | 57 |  | Injecting actual MCU failure is the only reliable method to verify the watchdog timeout triggers the correct safety response. Software simulation of the fault is insufficient for SIL 2 qualification. | verification, perception-subsystem, sil-2, session-701, idempotency:ver-sub026-701
VER-PERCEPTIONSUBSYSTEM-004 | Verify SUB-REQ-035: Mount vehicle on tilting test rig; rotate to 15.5 degrees at 5 degrees/second and record timestamp when tilt-hazard signal is asserted on Safety Monitor Processor I2C bus. Pass: signal asserted within 50ms of 15-degree threshold crossing for 5 trials in each axis. | verification-requirements | Perception Subsystem |  | Test | 57 |  | Controlled tilt rig test allows repeatable threshold measurement. I2C bus monitoring confirms the signal reaches the Safety Monitor Processor, not just the inclinometer output register. | verification, perception-subsystem, sil-2, session-702, idempotency:ver-sub035-702
VER-PERCEPTIONSUBSYSTEM-005 | Verify SUB-REQ-054: Apply vehicle tilt exceeding 15 degrees on test fixture and monitor tilt data publication rate and content to Safety Monitor Processor. Pass: data rate >= 20 Hz, includes tilt angle, axis identity, and confidence score, maintained for 30 s trial. | verification-requirements | Perception Subsystem |  | Test | 57 |  | SUB-REQ-054 feeds the safety halt chain; bus monitoring during tilt test validates both rate and content correctness simultaneously. | verification, perception-subsystem, session-705, idempotency:ver-sub054-705
VER-PERCEPTIONSUBSYSTEM-006 | Verify the Perception MCU processes all sensor inputs and emits an obstacle alert frame to the Safety Monitor Processor within 50ms of sensor data arrival at full 10Hz sensor update rate. | verification-requirements | Perception Subsystem |  | Test | 71 |  | SUB-REQ-024 specifies the 50ms processing and transmission latency at SIL-2; test must inject synthetic sensor frames at 10Hz and measure the CAN frame emission timestamp relative to sensor data arrival using a logic analyser, confirming 95th-percentile latency is ≤50ms over 1000 consecutive frames. | session-710, qc, perception-subsystem, sil-2, idempotency:ver-sub024-perception-latency-710
VER-PERCEPTIONSUBSYSTEM-007 | Verify the Side Proximity Sensor Pair detects objects within 0.5m laterally on each side and delivers a TTL-level alert signal to the Perception MCU within 30ms of intrusion. | verification-requirements | Perception Subsystem |  | Test | 71 |  | SUB-REQ-025 specifies 0.5m detection range and 30ms alert latency at SIL-2; test must position calibrated test objects at 0.5m, 0.45m, and 0.6m on each lateral side and confirm correct alert or no-alert states; latency from object placement to TTL signal assertion must be measured with an oscilloscope trigger. | session-710, qc, perception-subsystem, sil-2, idempotency:ver-sub025-proximity-sensors-710
VER-PERCEPTIONSUBSYSTEM-008 | Verify SUB-REQ-040: Connect a calibrated power supply to the Inclinometer Tilt Sensor Unit 3.3V input. Measure supply voltage at the sensor under nominal and maximum load conditions. Introduce a ±5% supply ripple at 1kHz and verify sensor output stability. Pass: supply voltage 3.3V ± 3% (3.20V–3.40V) under all load conditions; inclinometer output tilt reading stable ± 0.1° during ripple injection. Fail: voltage outside tolerance or tilt reading drift > 0.1° during ripple. | verification-requirements | Perception Subsystem |  | Test | 57 |  | SUB-REQ-040 is SIL-2; the inclinometer is the primary tip-over detection sensor (H-005). An out-of-tolerance power supply causes drift or false readings, either preventing safe-state activation or generating nuisance stops. The ±5% ripple test ensures resilience to DC-DC converter switching noise on the same rail. | session-711, validation, perception-subsystem, sil-2, idempotency:ver-sub040-inclinometer-power-711
VER-PERCEPTIONSUBSYSTEM-009 | Verify IFC-REQ-022: Position a 300mm × 300mm flat obstacle at 0.5m, 0.45m, and 0.6m on each lateral side of the vehicle. Pass: (a) alert asserted as TTL HIGH within 30ms of obstacle at 0.5m or 0.45m (≥ 20/20 placements per side); (b) no alert when obstacle at 0.6m (≥ 10/10 placements per side); (c) GPIO line returns LOW within 30ms of obstacle removal. Fail: any missed detection at ≤ 0.5m, any false positive at 0.6m, or delayed deassertion. | verification-requirements | Perception Subsystem |  | Test | 57 |  | IFC-REQ-022 specifies the lateral obstacle GPIO interface feeding the collision prevention chain (H-002 SIL-2). Testing at 0.45m confirms margin below the 0.5m threshold; testing at 0.6m confirms no over-reaching false positives that would generate nuisance stops in doorway-width corridors. | session-711, validation, perception-subsystem, sil-2, idempotency:ver-ifc022-proximity-gpio-711
VER-PERCEPTIONSUBSYSTEM-010 | Verify IFC-REQ-021: Scan I2C bus at 400 kHz (400 kHz I2C mode per UM10204) and enumerate all three depth sensor addresses. Drive each sensor with a ranging command and verify response timing. Pass: (a) all three sensors enumerate at correct addresses; (b) ranging responses received within 10ms of command; (c) no bus contention (SDA/SCL monitored with logic analyser) across 1000 ranging cycles. Fail: any address absent, latency > 10ms, or SDA collision detected. | verification-requirements | Perception Subsystem |  | Test | 57 |  | IFC-REQ-021 specifies the I2C interface for the Forward Depth Sensor Array feeding H-002 collision prevention. Bus contention or missing sensors would blind the obstacle detection system at operating speed. Testing all three addresses ensures no single device can disappear silently from the bus. | session-711, validation, perception-subsystem, sil-2, idempotency:ver-ifc021-depth-sensor-i2c-711
VER-POWERSUBSYSTEM-001 | The Battery Management System thermal fault response SHALL be verified by forcing a cell temperature sensor input above 60°C via calibrated resistance substitution and measuring time to 48V bus de-energisation. Pass: bus de-energised within 200ms and CAN fault code broadcast within 200ms on ≥5/5 test runs. | verification-requirements | Power Subsystem |  | Test | 71 |  | Direct hardware-in-the-loop verification of SIL-3 thermal protection function. Resistance substitution simulates worst-case NTC sensor input without requiring actual overheating of cells. 5 consecutive pass runs establishes repeatability for safety case submission. | verification, power-subsystem, sil-3, session-699, idempotency:ver-sub013-bms-thermal-cutoff-699
VER-POWERSUBSYSTEM-002 | The BMS-to-Safety Monitor Processor GPIO fault interface SHALL be verified by injecting a 55°C temperature threshold crossing and measuring: (a) GPIO fault line transitions low within 50ms; (b) transition is independent of CAN bus availability. Pass: both criteria met on ≥5/5 runs with CAN connected and ≥5/5 runs with CAN disconnected. | verification-requirements | Power Subsystem |  | Test | 71 |  | Independence of the fault line from CAN is the critical safety property. Test must be run both with and without CAN to demonstrate independence, as required for IEC 61508 (Functional safety of E/E/PE safety-related systems) SIL-3 diagnostic coverage. | verification, power-subsystem, sil-3, session-699, idempotency:ver-ifc016-bms-fault-line-699
VER-POWERSUBSYSTEM-003 | Verify SUB-REQ-071: Connect the buggy to both a 230V 50Hz supply and (separately) a 120V 60Hz supply at 20% SoC. Measure time to reach 100% SoC, confirming CC-CV profile is applied and charge terminates automatically. Pass if charge completes within 4 hours on both supply voltages, SoC measurement accuracy ±2%, and trickle maintenance engages on completion. | verification-requirements | Power Subsystem |  | Test | 57 |  | Full charge test from 20% SOC validates the charge profile and timing under both mains standards. ±2% SOC accuracy is achievable with a calibrated BMS coulomb counter and confirms the test is deterministic. | verification, power-subsystem, session-707, idempotency:ver-sub-req-071-707, tech-author-session-719
VER-POWERSUBSYSTEM-004 | Verify the Lithium Iron Phosphate Battery Pack delivers a minimum of 4 hours continuous operation at rated system load: (a) charge pack to 100% SoC at 25°C, (b) discharge at combined rated system load (drive motors at 3 km/h on flat surface, BCI processing active, HMI and Safety Subsystem energised), (c) record elapsed time from start of discharge to first alarm at 10% SoC. Repeat at 40°C ambient. Pass: elapsed time ≥ 4 hours at both 25°C and 40°C. Fail: either condition below 4 hours. | verification-requirements | Power Subsystem |  | Test | 57 |  | SUB-REQ-014 specifies the 4-hour minimum runtime; test must discharge the battery pack at full combined system load (drive, BCI processing, HMI, safety systems) from 100% to 10% SoC and record elapsed time, at 25°C and at 40°C boundary condition. | session-710, qc, power-subsystem, sil-3, idempotency:ver-sub014-battery-runtime-710, tech-author-session-719
VER-POWERSUBSYSTEM-005 | Verify the DC-DC Converter Array maintains each output rail (12V, 5V, 3.3V) within ±5% of nominal voltage and with output ripple not exceeding 50mV peak-to-peak across 10% to 100% rated load range. | verification-requirements | Power Subsystem |  | Test | 71 |  | SUB-REQ-015 specifies rail regulation and ripple limits; measurement must be taken with a calibrated oscilloscope at each load step increment across both temperature extremes (0°C and 40°C) to confirm worst-case regulation is within specification for all downstream electronics. | session-710, qc, power-subsystem, sil-3, idempotency:ver-sub015-dcdc-regulation-710
VER-POWERSUBSYSTEM-006 | Verify the Charge Controller charges the battery pack from 20% to 100% state of charge within 3 hours when connected to a 240V AC facility supply. | verification-requirements | Power Subsystem |  | Test | 71 |  | SUB-REQ-016 specifies the 3-hour fast-charge bound; test must time three complete charge cycles from 20% SoC on a calibrated supply simulator at 240V nominal, recording SoC vs time, cell temperatures, and charger current profile to confirm charge completion and thermal safety. | session-710, qc, power-subsystem, sil-3, idempotency:ver-sub016-charge-time-710
VER-POWERSUBSYSTEM-007 | Verify the Battery Management System transmits pack voltage, individual cell voltages, SoC, temperature summary, and pack current to the Main Application Processor via CAN 2.0B at 250 kbit/s at 100ms intervals. | verification-requirements | Power Subsystem |  | Test | 57 |  | IFC-REQ-017 specifies the critical battery monitoring CAN bus interface at SIL-3; test must capture 1000 CAN frames with a protocol analyser, decode each against the specified message format, confirm all five data fields are present and within physical ranges, verify 250 kbit/s bus speed, and confirm inter-frame interval is 100ms ±5ms. | session-710, qc, power-subsystem, sil-3, idempotency:ver-ifc017-bms-can-710
VER-POWERSUBSYSTEM-008 | Verify IFC-REQ-002: Connect the vehicle to a 230V 50Hz facility charging dock and measure galvanic isolation resistance between vehicle chassis and mains supply. Pass: (a) isolation resistance ≥ 2MΩ at 500V DC (per IEC 60601-1 measurement method); (b) charge current flows within 60 seconds of dock connection; (c) charge terminates automatically at 100% SoC with charger LED indicating complete. Fail: isolation resistance < 2MΩ, no charging within 60s, or no auto-termination. | verification-requirements | Power Subsystem |  | Test | 57 |  | IFC-REQ-002 charging interface galvanic isolation is a safety requirement protecting a non-ambulatory user from mains-frequency shock while docked in a bed or care facility. IEC 60601-1 mandates 2MΩ isolation for patient-adjacent devices. | session-711, validation, power-subsystem, idempotency:ver-ifc002-charging-iso-711
VER-POWERSUBSYSTEM-009 | Verify SUB-REQ-036: Connect the charge controller to a 230V 50Hz supply and then a 120V 60Hz supply. Measure input power acceptance and output to battery. Pass: (a) charging initiates within 60s at 230V 50Hz; (b) charging initiates within 60s at 120V 60Hz; (c) no fault codes generated at either input; (d) BMS reports valid charge current at both supply voltages. Fail: failure to accept either supply or fault code generated. | verification-requirements | Power Subsystem |  | Test | 57 |  | SUB-REQ-036 covers universal AC input for global facility deployment; a charge controller that rejects one mains standard would prevent charging at overseas facilities and require a hardware variant. The test confirms both operating ranges without hardware change. | session-711, validation, power-subsystem, idempotency:ver-sub036-charge-ctrl-ac-711
VER-POWERSUBSYSTEM-010 | Verify SUB-REQ-055: Connect the Power Subsystem to a 230V 50Hz AC supply at 20% SoC via a calibrated current probe on the mains lead. (a) Measure peak inrush current during first 100ms of connection across 5 trials — Pass: peak ≤16 A on all 5 trials. (b) Monitor cell temperature via thermocouple array at 1Hz during full charge cycle — Pass: no cell exceeds 45°C. (c) Time charge completion from 20% to 100% SoC — Pass: completion within 6 hours. Fail: inrush >16 A, cell temperature >45°C, or charge time >6 hours on any trial. | verification-requirements | Power Subsystem |  | Test | 57 |  | SUB-REQ-055 derives from the electrical safety mandate for 230V facility infrastructure (16A circuit breaker protection). Inrush >16A trips facility breakers and disrupts other equipment. Cell temperature >45°C accelerates electrolyte degradation and raises H-003 thermal runaway risk. End-to-end measurement at the mains lead is necessary because inrush depends on the complete circuit including charger topology and cable impedance. | session-713, validation, power-subsystem, idempotency:ver-sub055-charging-713
VER-SAFETYSUBSYSTEM-001 | Verify SUB-REQ-001: Inject a simulated main processor fault (power rail pull-down to 0V) while Safety Monitor Processor is executing safety state machine. Verify via logic analyser that safety state machine continues to execute at 200Hz for ≥5 seconds post-fault, and that relay de-energise command is issued within 500ms of heartbeat loss detection. Pass: no execution gap >6ms in safety loop; relay command issued; all outputs set to safe state. Fail: execution halt, missed heartbeat detection, or relay remains energised. | verification-requirements | Safety Subsystem |  | Test | 57 |  | Physical fault injection is the only method that truly verifies processor independence — software simulation of the fault on the same processor cannot distinguish independent operation from shared resource operation. | verification, safety-subsystem, session-697, idempotency:ver-sub001-smp-independence-697
VER-SAFETYSUBSYSTEM-002 | Verify SUB-REQ-002: Apply de-energise command to Motor Power Isolation Relay under 4 load conditions (0A, 30A, 60A steady, 200A pulse). Measure with oscilloscope the time from control-line falling edge to relay contact open confirmed by feedback contact. Pass: all measurements ≤20ms. Fail: any measurement >20ms, or relay fails to open within 50ms under any load condition. | verification-requirements | Safety Subsystem |  | Test | 57 |  | Relay switching time is load-dependent — the test must cover the full rated range including peak inrush current to validate the 20ms requirement is met at worst case. | verification, safety-subsystem, session-697, idempotency:ver-sub002-relay-response-697
VER-SAFETYSUBSYSTEM-003 | Verify SUB-REQ-003: Replay 50 CHB-MIT annotated seizure EEG segments through the Seizure Detection Module and measure time-to-detection flag from annotated onset. Also replay 8 hours of normal EEG operation and count false positive seizure flags. Pass: ≥95% of seizure segments detected within 150ms; false positive count ≤1 across full 8-hour replay. Fail: detection latency >150ms for >5% of segments, or false positive count >1. | verification-requirements | Safety Subsystem |  | Test | 57 |  | Open benchmark dataset (CHB-MIT) allows reproducible and auditable verification. The 95%/150ms criterion aligns with the clinical sensitivity requirement — some ambiguous seizure onset boundaries account for the 5% allowance. | verification, safety-subsystem, session-697, idempotency:ver-sub003-seizure-detect-697
VER-SAFETYSUBSYSTEM-004 | Verify IFC-REQ-007: Inject a simulated main processor heartbeat halt (hold GPIO low). Measure on oscilloscope from last valid heartbeat to Safety Monitor Processor E-stop output assertion. Separately, verify no data payload is transmitted on the heartbeat line using a protocol analyser. Pass: E-stop asserted within 500ms of last heartbeat; no data payload detected on heartbeat GPIO. Fail: E-stop delayed >500ms or data content detected. | verification-requirements | Safety Subsystem |  | Test | 57 |  | Integration test verifies that the heartbeat interface contract is respected end-to-end — both the timing (500ms watchdog window) and the structural independence (no data sharing on the heartbeat wire). | verification, safety-subsystem, session-697, idempotency:ver-ifc007-heartbeat-697
VER-SAFETYSUBSYSTEM-005 | Verify IFC-REQ-008: With relay energised, open (disable) one control channel at a time. Verify relay de-energises within 20ms for each single-channel failure. Also verify relay does not energise when both channels are driven independently by separate fault-injected outputs. Pass: relay opens on either single-channel fault; relay does not energise with mismatched channels. Fail: relay remains closed on single-channel fault. | verification-requirements | Safety Subsystem |  | Test | 57 |  | Verifies the 1oo2 architecture's fail-safe property — the test must exercise each channel independently because a symmetric test could mask a shared failure mode in the dual-channel driver circuit. | verification, safety-subsystem, session-697, idempotency:ver-ifc008-relay-dual-697
VER-SAFETYSUBSYSTEM-006 | Verify SYS-REQ-002 end-to-end: In a fully integrated vehicle at maximum speed (6 km/h), inject each of the 5 safety triggers (BCI signal loss, seizure EEG replay, tilt >15°, manual E-stop button, processor heartbeat loss) and measure from trigger event to vehicle coming to rest (wheels stopped). Pass: vehicle comes to rest within 200ms plus brake stopping distance for all 5 triggers. Fail: any trigger results in continued motion beyond 200ms relay actuation time. | verification-requirements | Safety Subsystem |  | Test | 57 |  | System-level integration test validates the complete safety chain — detection, relay actuation, and mechanical braking — under realistic conditions. Component-level tests cannot detect emergent timing failures introduced by integration. | verification, safety-subsystem, session-697, idempotency:ver-sys002-estop-e2e-697
VER-SAFETYSUBSYSTEM-007 | The Inclinometer Tilt Sensor Unit SHALL be verified by mounting the buggy on a calibrated tilt table, incrementally inclining from 0° to 20° in 1° steps at 1°/s, and confirming: (a) angle readings within ±0.5° of reference at each step, (b) 100Hz output rate continuous across full range, (c) E-stop command issued by Safety Monitor Processor within 200ms of crossing 15° sustained for 200ms. | verification-requirements | Safety Subsystem |  | Test | 86 |  | Tilt accuracy and E-stop trigger thresholds are safety-critical (SIL 2, H-005 tip-over). End-to-end test on a tilt table is the only method that validates sensor accuracy, update rate, and the Safety Monitor Processor response as an integrated chain. Analysis alone cannot verify the 200ms propagation latency through the hardware interface. | verification, safety-subsystem, session-698, idempotency:ver-inclinometer-tilt-698
VER-SAFETYSUBSYSTEM-008 | The Manual Emergency Stop Button hardwire circuit SHALL be verified by: (a) measuring contact-to-relay-coil de-energisation time using an oscilloscope across 10 actuations with motor at rated load, confirming all measurements ≤5ms, (b) inspecting the wiring schematic to confirm no software-mediated path in the disconnect circuit. | verification-requirements | Safety Subsystem |  | Test | 86 |  | SUB-REQ-005 specifies a hardware-only disconnect path to ensure the E-stop works even during software failure. Oscilloscope measurement is required because the 5ms budget is too short to verify by any other method, and schematic inspection confirms the hardwire bypass requirement. Safety-critical for IEC 61508 SIL 3 independence claim. | verification, safety-subsystem, session-698, idempotency:ver-estop-hardwire-698
VER-SAFETYSUBSYSTEM-009 | The Safety Subsystem E-stop state machine SHALL be verified by injecting each of the five defined safety triggers (BCI signal loss, seizure flag, tilt threshold, manual E-stop, heartbeat failure) individually and in combination, confirming: relay de-energisation, brake command, and HMI alert all activated within 200ms of trigger, measured from trigger signal assertion to final output using a logic analyser. | verification-requirements | Safety Subsystem |  | Test | 86 |  | SUB-REQ-006 is the primary safety requirement covering all E-stop trigger paths (SIL 3). Fault injection test is the IEC 61508-mandated verification method for safety function completeness — analysis cannot confirm the combined-trigger behaviour or measure real latency through the hardware interrupt chain. | verification, safety-subsystem, session-698, idempotency:ver-estop-state-machine-698
VER-SAFETYSUBSYSTEM-010 | The Safety Monitor Processor heartbeat watchdog SHALL be verified by: (a) halting the main application processor and confirming E-stop triggers within 600ms (500ms watchdog timeout + 100ms response margin), (b) injecting heartbeat period jitter at +20% and +21% of nominal and confirming E-stop triggers only at the 21% case, (c) repeating 50 times to confirm deterministic behaviour. | verification-requirements | Safety Subsystem |  | Test | 86 |  | SUB-REQ-007 implements the software-fault detection path of the SIL 3 safety chain. Boundary condition testing at ±20% heartbeat deviation is required to verify the threshold is correctly implemented — a simple timeout test alone does not cover the period-deviation trigger condition. | verification, safety-subsystem, session-698, idempotency:ver-heartbeat-watchdog-698
VER-SAFETYSUBSYSTEM-011 | The emergency BLE iBeacon interface SHALL be verified by: (a) triggering each of five E-stop conditions and measuring time from trigger to iBeacon advertisement transmission at facility BLE gateway, confirming < 500ms in 50 consecutive tests; (b) confirming advertisement contains correct device ID, alert type, and GPS coordinates by sniffing the BLE advertisement payload with a protocol analyser. | verification-requirements | Safety Subsystem |  | Test | 86 |  | IFC-REQ-003 implements the attendant alerting function required by STK-REQ-005. The 500ms latency requirement is safety-critical — delayed alerts mean the carer cannot respond in time to assist a seizing or stranded user. End-to-end test at the facility BLE gateway is required because relay latency depends on gateway polling interval, not just buggy transmission time. | verification, safety-subsystem, session-698, idempotency:ver-ble-emergency-alert-698
VER-SAFETYSUBSYSTEM-012 | The SPI interface between Inclinometer Tilt Sensor Unit and Safety Monitor Processor SHALL be verified by: (a) capturing 1000 consecutive SPI transactions at 1MHz with a logic analyser, confirming 6-byte frame format, roll/pitch/timestamp fields, and CRC validity in all frames; (b) injecting a corrupted CRC and confirming the Safety Monitor Processor discards the frame and logs a sensor fault within one SPI cycle (1ms). | verification-requirements | Safety Subsystem |  | Test | 86 |  | IFC-REQ-009 carries safety-critical tilt data. CRC validation is the sole defence against corrupted angle readings that could prevent an E-stop during tip-over (H-005, SIL 2). Logic analyser capture is required to verify the frame format and CRC implementation before integration testing. | verification, safety-subsystem, session-698, idempotency:ver-spi-tilt-smp-698
VER-SAFETYSUBSYSTEM-013 | Verify SUB-REQ-043: Inspect electronic assembly against design documentation. Confirm Safety Monitor Processor PCB is mechanically separate from Main Application Processor PCB, has independent connector to DC-DC Converter Array power rail, and has no shared power nets with the Main Application Processor on the backplane. | verification-requirements | Safety Subsystem |  | Inspection | 57 |  | Physical separation is an assembly attribute verifiable by inspection against engineering drawings. Power independence must be confirmed at net level via schematic review. | verification, safety-subsystem, sil-3, session-702, idempotency:ver-sub043-702
VER-SAFETYSUBSYSTEM-014 | Verify SYS-REQ-020: review EEG headset manufacturer's biological safety evaluation report against ISO 10993-1. Confirm biocompatibility testing covers cytotoxicity (ISO 10993-5), sensitisation (ISO 10993-10), and skin irritation (ISO 10993-23) for all skin-contacting electrode and headset materials. Pass: manufacturer provides valid ISO 10993-1 conformance certificate with no Category I or Category II adverse findings for prolonged skin contact. | verification-requirements | Safety Subsystem |  | Inspection | 57 |  | ISO 10993-1 compliance for skin-contact medical devices is verified by documentation inspection of the manufacturer biological safety evaluation, not by independent lab testing. Pass/fail criterion anchors to specific test standard clause levels. | verification, biocompatibility, safety, session-703, idempotency:ver-sys020-703
VER-SAFETYSUBSYSTEM-015 | Verify SUB-REQ-051: Simulate facility emergency system halt command via 2.4 GHz wireless interface and measure elapsed time from signal receipt to motor isolation relay actuation and electromagnetic brake engagement. Pass: both events complete within 150 ms across 20 consecutive trials at operating temperature. | verification-requirements | Safety Subsystem |  | Test | 57 |  | SUB-REQ-051 is safety-critical (SIL-3 chain); timed integration test with actual RF link validates both the hardware path and software response, not just simulation. | verification, safety-subsystem, session-705, idempotency:ver-sub051-705
VER-SAFETYSUBSYSTEM-016 | Verify SUB-REQ-057: Halt MAP watchdog refresh and measure elapsed time to SAFE_STOP assertion and non-volatile fault log write. Pass: SAFE_STOP asserted within 200 ms, fault record present in NV storage, across 10 trials. | verification-requirements | Safety Subsystem |  | Test | 71 |  | SUB-REQ-057 is a SIL-2 failover requirement; direct watchdog injection on test hardware is the only reliable method to verify the timing chain without simulation. | verification, safety-subsystem, session-705, sil-2, idempotency:ver-sub057-705
VER-SAFETYSUBSYSTEM-017 | Verify SUB-REQ-060: Using a test fixture that monitors relay coil continuity, open each NC contact channel independently (channel A, then channel B) and confirm Motor Power Isolation Relay de-energises within 20ms measured from contact open to relay drop-out on oscilloscope. Pass criterion: de-energisation within 20ms for both channels independently, zero relay hold when either channel is open. | verification-requirements | Safety Subsystem |  | Test | 57 |  | Dual-channel verification must be done per-channel to confirm independence; a combined test that opens both channels simultaneously does not prove single-channel sufficiency, which is the safety claim. | verification, safety-subsystem, session-706, sil-3, idempotency:ver-sub060-estop-dual-706
VER-SAFETYSUBSYSTEM-018 | Verify SUB-REQ-061: In a controlled test environment, inject a simulated MAP watchdog timeout by halting MAP firmware execution, and verify that the Safety Monitor Processor asserts Emergency Stop state and zeros velocity command within 250ms (200ms watchdog + 50ms propagation) with no dependency on MAP recovery. Pass criterion: velocity = 0 within 250ms across 10 injected fault events, and vehicle does not resume motion without manual reset. | verification-requirements | Safety Subsystem |  | Test | 57 |  | A test that waits for watchdog timeout and measures SMP response independently confirms the failover path operates without MAP participation, which is the safety claim. | verification, safety-subsystem, session-706, sil-3, idempotency:ver-sub061-map-failover-706
VER-SAFETYSUBSYSTEM-019 | Verify SUB-REQ-063: Review certification documentation from an accredited IEC 61508 assessor covering Safety Subsystem hardware fault tolerance (HFT), systematic capability SC3, and software development process documentation (lifecycle plan, V&V evidence). Pass criterion: certificate issued by a Notified Body or TUV-equivalent confirming SIL-3 suitability and no open safety defects. | verification-requirements | Safety Subsystem |  | Inspection | 57 |  | SIL-3 certification requires third-party evidence; this cannot be verified by testing alone. Document inspection of the assessment report is the appropriate verification method per IEC 61508-1 clause 6.2. | verification, safety-subsystem, session-706, sil-3, idempotency:ver-sub063-sil3-cert-706
VER-SAFETYSUBSYSTEM-020 | Verify SUB-REQ-070: Place the buggy on a tilt platform and gradually increase pitch to 12 degrees (confirm pre-warning signal issued to Safety Monitor), then continue to 15 degrees (confirm brake application and motion inhibit). Measure time from 15-degree crossing to brake confirmed. Pass if pre-warning issued at 12±0.5 degrees, brake applied within 50ms of 15-degree crossing, and motion inhibited within 100ms, on 5 trials each axis. | verification-requirements | Safety Subsystem |  | Test | 57 |  | Tilt-platform test directly validates the two-threshold detection scheme and timing. 5 trials per axis (pitch and roll) verify the requirement is met across the full mechanical envelope. The ±0.5 degree tolerance is achievable with the specified sensor at 20Hz sampling rate. | verification, safety-subsystem, sil-2, session-707, idempotency:ver-sub-req-070-707
VER-SAFETYSUBSYSTEM-021 | The MAP watchdog failsafe mechanism SHALL be verified by injecting a simulated MAP watchdog timeout (halt firmware refresh) and measuring time from last heartbeat to SAFE_STOP assertion and brake engagement; the total latency SHALL be ≤200ms measured on oscilloscope at the relay coil. | verification-requirements | Safety Subsystem |  | Test | 86 |  | Provides specific measurable test procedure for SYS-REQ-023. Oscilloscope measurement at relay coil provides unambiguous physical evidence of brake engagement timing, avoiding software timing artifacts. This test is equivalent to VER-REQ-046 (SUB-REQ-057) but at system integration level, confirming end-to-end latency including SMP decision time and relay actuation. | verification, safety, session-708, qc, idempotency:ver-map-watchdog-failover-708
VER-SAFETYSUBSYSTEM-022 | Verify the Safety Subsystem Safety Monitor Processor is developed and validated in accordance with IEC 61508-3 at SIL 3, with documented V-model lifecycle artefacts. | verification-requirements | Safety Subsystem |  | Inspection | 71 |  | SUB-REQ-059 mandates IEC 61508-3 (Functional Safety of E/E/PE safety-related systems, Part 3: Software requirements) SIL-3 development; inspection must confirm existence of hazard and risk analysis, software safety requirements specification, SIL-3 architecture description, and independent verification report, with each artefact referencing IEC 61508-3 clauses. | session-710, qc, safety-subsystem, sil-3, standards, idempotency:ver-sub059-smp-iec61508-710
VER-SAFETYSUBSYSTEM-023 | Verify the Seizure Detection Module updates the shared memory region on the Safety Monitor Processor at 50Hz with a 1-byte status word and that the Safety Monitor Processor reads this region within 10ms of each update. | verification-requirements | Safety Subsystem |  | Test | 57 |  | IFC-REQ-010 defines a safety-critical shared memory IPC path between SDM and SMP; test must inject SDM updates at 50Hz with controlled timestamps and measure the SMP read latency using RTOS task tracing, confirming 99th-percentile read latency ≤10ms and mutex contention does not cause missed frames over a 5-minute stress test. | session-710, qc, safety-subsystem, idempotency:ver-ifc010-sdm-smp-shmem-710
VER-VEHICLEPLATFORM-001 | Verify SUB-REQ-027: load test the Chassis Frame per ISO 7176-8 static load procedure — apply 150kg distributed load via seat mounting points, hold for 5 minutes, inspect for permanent deformation. Apply 3g vertical impulse via drop test fixture. Pass: zero permanent deformation after static test; no fracture or crack after dynamic test. | verification-requirements | Vehicle Platform |  | Test | 57 |  | ISO 7176-8 defines the standard structural test procedure for powered wheelchairs and is the accepted verification method for regulatory submission. | verification, vehicle-platform, session-701, idempotency:ver-sub027-701
VER-VEHICLEPLATFORM-002 | Verify SUB-REQ-049: perform minimum turning radius test per ISO 7176-3 test method. Mark vehicle centre on test floor, execute full 360-degree turning manoeuvre, measure outer wheel track swept radius. Verify navigation through a 900mm doorway simulation with 150mm marker barriers on each side. Pass: swept radius <= 600mm, no contact with barriers in 5 trials. | verification-requirements | Vehicle Platform |  | Test | 57 |  | Validates the turning radius budget against accessible doorway navigation requirement derived from SYS-REQ-010 and BS 8300 compliance. | verification, vehicle-platform, session-703, idempotency:ver-sub049-703
VER-VEHICLEPLATFORM-003 | The Electronics Bay IP54 ingress protection rating SHALL be verified by performing the IEC 60529 (Degrees of protection provided by enclosures) test method 13.4 dust wire probe and 14.2.4 water splash tests, plus fastener torque and gasket continuity inspection. Pass: no dust ingress at IP5X probe; no functional degradation after water splash; all bay fasteners at specified torque. | verification-requirements | Vehicle Platform |  | Test | 71 |  | IP54 is a defined test per IEC 60529; standardised test procedures provide objective pass/fail evidence for the housing protection claim. | verification, vehicle-platform, session-706, idempotency:ver-sys021-ip54-706
VER-VEHICLEPLATFORM-004 | The Chassis Frame SHALL be verified by static load testing per ISO 7176-8: apply 1.5× (maximum user mass 130kg + electronics bay 15kg) = 217.5kg to the seat mounting points and hold for 60 seconds; no permanent deformation >1mm shall be measured at any structural member. Impact test: drop a 7kg mass from 1m height onto each corner of the chassis platform and confirm no fracture or deformation of safety-critical welds. | verification-requirements | Vehicle Platform |  | Test | 86 |  | Directly verifies SUB-REQ-078 structural compliance claims. Static load test replicates worst-case seated user mass per ISO 7176-8. Impact test replicates doorway kerb and ramp transitions at care facility speeds. Both tests use ISO 7176-8 test fixtures and measurement procedures to ensure results are comparable with regulatory submission requirements. | verification, vehicle-platform, session-708, qc, idempotency:ver-chassis-iso7176-compliance-708
VER-VEHICLEPLATFORM-005 | Verify the Electronics Bay maintains internal temperature below 70°C for all mounted electronics while dissipating up to 40W at 40°C ambient with vehicle stationary and no forced airflow. | verification-requirements | Vehicle Platform |  | Test | 71 |  | SUB-REQ-028 specifies the 70°C internal temperature limit under worst-case thermal load; thermal test must run the vehicle system at full electrical load in a 40°C environmental chamber for 4 hours and confirm no internal thermocouple reading exceeds 70°C at any point during steady-state operation. | session-710, qc, vehicle-platform, idempotency:ver-sub028-ebay-thermal-710
VER-VEHICLEPLATFORM-006 | Verify the Electronics Bay provides IP54 environmental protection per IEC 60529 (Degrees of protection provided by enclosures): (a) expose assembled unit to IP5X dust chamber per IEC 60529 Clause 13.4 for 8 hours; (b) expose to IP54 water spray per IEC 60529 Clause 14.2.4 for 5 minutes at all orientations; (c) open and re-close access panel 100 times; (d) repeat dust and water tests. Pass: no dust ingress visible on internal surfaces after step (a); no water ingress after step (b); IP54 rating maintained after 100 seal cycles in step (d). Fail: any ingress. | verification-requirements | Vehicle Platform |  | Test | 57 |  | SUB-REQ-050 mandates IP54 per IEC 60529 (Degrees of protection provided by enclosures) and 100 access cycle seal durability; testing must include: dust chamber exposure per IEC 60529 IP5X procedure, water spray test per IP54X procedure, and repeated seal closure/opening for 100 cycles followed by repeat IP54 tests confirming no seal degradation. | session-710, qc, vehicle-platform, idempotency:ver-sub050-ip54-710
VER-VEHICLEPLATFORM-007 | Verify SUB-REQ-074: Inspect the chassis frame material certificate and dimensional report. Pass: (a) material certificate confirms aluminium alloy 6061-T6 or equivalent per ISO 209 (Wrought aluminium and aluminium alloys); (b) cross-section dimensional report confirms wall thickness ≥ 3mm at all primary load-bearing members; (c) static load test to 1.5× rated occupant load (180kg) with no permanent deformation > 1mm. Fail: any non-conforming material, wall thickness < 3mm, or deformation > 1mm. | verification-requirements | Vehicle Platform |  | Test | 57 |  | SUB-REQ-074 is SIL-2; chassis structural failure under load could cause tip-over (H-005) or entrapment of the non-ambulatory user. The 1.5× static overload proof test is the minimum acceptance criterion per EN 12184 (electrically powered wheelchairs and scooters) design verification. | session-711, validation, vehicle-platform, sil-2, idempotency:ver-sub074-chassis-material-711
VER-VEHICLEPLATFORM-008 | Verify SUB-REQ-072: Connect a service laptop to the Electronics Bay USB-C 3.1 Gen 1 port and execute the automated diagnostic test suite. Pass: (a) enumeration completes within 10 seconds of cable connection; (b) diagnostic API returns responses for all five test categories (motor, brake, sensor, battery, BCI pipeline) within 20 minutes; (c) test suite returns PASS/FAIL result for each subsystem with numeric values and fault codes. Fail: any test category fails to return result or API returns null. | verification-requirements | Vehicle Platform |  | Test | 57 |  | SUB-REQ-072 covers the Weekly Maintenance scenario diagnostic port. The 20-minute test suite limit maps to the 20 min/buggy maintenance window for a 4-buggy fleet. Failure to enumerate or incomplete API response blocks the maintenance workflow and could allow undetected faults to enter service. | session-711, validation, vehicle-platform, maintenance, idempotency:ver-sub072-diag-usbc-711
VER-VEHICLEPLATFORM-009 | The Electronics Bay IP54, thermal, and access-cycle durability SHALL be verified by: (a) IEC 60529 dust probe and 10-minute water spray test — Pass: no wire penetration, no functional degradation; (b) 4-hour operation at 55°C ambient at maximum component load — Pass: processor inlet temperature ≤55°C, no thermal shutdown; (c) 100 maintenance access cycles followed by re-test of (a) — Pass: IP54 rating maintained. | verification-requirements | Vehicle Platform |  | Test | 86 |  | SUB-REQ-075 is SIL-1 rated and directly supports SUB-REQ-050 and SYS-REQ-021. Hospital care facilities use wet floor cleaning with mop spray routinely. Thermal management failure above 55°C causes uncontrolled processor shutdown, voiding all real-time safety functions. The 100-cycle maintenance endurance test ensures IP rating persists through the 5-year service life at quarterly inspection intervals as required by IEC 60601-1 clause 11.6.5. | session-713, validation, vehicle-platform, sil-1, idempotency:ver-sub075-ebay-ip54-713
VER-VEHICLEPLATFORM-010 | Verify SUB-REQ-029: Using a calibrated load test rig with adjustable torso mannequins at the extremes (40 kg/1400mm stature and 120 kg/1900mm stature): (a) Confirm backrest adjusts ≥15 degrees of recline at both occupant configurations. (b) Apply 3g forward deceleration pulse via sled test and measure lap belt restraint force — Pass: belt remains engaged, forward displacement of 50th-percentile mannequin torso ≤50mm from seated position. (c) Confirm 4-point harness buckle engages audibly and cannot be released by a single-finger pull force <20N. Fail: range below 15°, torso displacement >50mm, or buckle release <20N. | verification-requirements | Vehicle Platform |  | Test | 57 |  | SUB-REQ-029 is the primary occupant restraint requirement for non-ambulatory users who cannot self-brace during emergency braking. The 3g pulse corresponds to the maximum deceleration during emergency braking at 6 km/h on a non-compliant surface. Extremes of the occupant range must be tested because harness geometry is mass/stature-dependent. EN 12184 (electric wheelchairs) requires dynamic crash testing of restraint systems. | session-713, validation, vehicle-platform, idempotency:ver-sub029-seat-restraint-713
VER-VEHICLEPLATFORM-011 | The Electronics Bay chassis mounting interface SHALL be verified by: (a) axial pull-out test at 50N/s ramp on four M6 fasteners with self-locking nuts and rubber isolators — Pass: no fastener pulls out below 500N, no isolator de-bonds below 200N; (b) 3g 11ms half-sine shock test — Pass: no isolator failure, no fastener torque loss >10%. | verification-requirements | Vehicle Platform |  | Test | 71 |  | IFC-REQ-024 specifies the electronics bay mechanical attachment — failure allows the bay (2 kg) to impact the user or obstruct drive components during an emergency stop deceleration event. 500N pull-out strength provides 8x margin over the 3g inertia load at 2 kg mass. Rubber isolators prevent resonance damage to processor PCBs rated for 2g sinusoidal vibration. | session-713, validation, vehicle-platform, idempotency:ver-ifc024-ebay-fasteners-713
VER-VEHICLEPLATFORM-012 | Verify IFC-REQ-031: (a) Test quick-release mechanism on all four wheels using a calibrated force gauge — apply force in the release direction, measure peak force for removal. Pass: all wheels release below 15N. (b) Partially engage the mechanism by stopping at mid-travel and apply 200N axial load — Pass: mechanical interlock prevents partial engagement; wheel fully locks or fully releases, no intermediate state. (c) Install all four wheels and verify positive engagement (audible click, no rotational play under 50N lateral load). Fail: release force >15N, partial engagement holds under load, or wheel rotates under lateral load. | verification-requirements | Vehicle Platform |  | Test | 57 |  | IFC-REQ-031 governs wheel removal during maintenance, which occurs weekly. If release force exceeds 15N, technicians apply excess force risking injury or mechanism damage. The interlock test is safety-critical: a partially engaged wheel could detach during operation causing H-005 (tip-over) or H-002 (collision). Lateral load test confirms the locking mechanism resists road forces during normal navigation. | session-713, validation, vehicle-platform, maintenance, idempotency:ver-ifc031-wheel-release-713

## Trace Links
Source Ref | Target Ref | Link Type | Description | Rationale | Created
SUB-POWERSUBSYSTEM-006 | SUB-POWERSUBSYSTEM-014 | derives | SUB-REQ-055 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-055: it specifies the system fault response when the nominal function defined in SUB-REQ-055 is not achieved within the specified performance threshold. | 2026-03-31T10:14:26.310Z
SUB-PERCEPTIONSUBSYSTEM-007 | SUB-PERCEPTIONSUBSYSTEM-012 | derives | SUB-REQ-054 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-054: it specifies the system fault response when the nominal function defined in SUB-REQ-054 is not achieved within the specified performance threshold. | 2026-03-31T10:14:25.493Z
SUB-POWERSUBSYSTEM-005 | SUB-POWERSUBSYSTEM-012 | derives | SUB-REQ-036 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-036: it specifies the system fault response when the nominal function defined in SUB-REQ-036 is not achieved within the specified performance threshold. | 2026-03-31T10:14:24.692Z
SUB-PERCEPTIONSUBSYSTEM-005 | SUB-PERCEPTIONSUBSYSTEM-011 | derives | SUB-REQ-035 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-035: it specifies the system fault response when the nominal function defined in SUB-REQ-035 is not achieved within the specified performance threshold. | 2026-03-31T10:14:23.906Z
SUB-POWERSUBSYSTEM-003 | SUB-POWERSUBSYSTEM-010 | derives | SUB-REQ-015 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-015: it specifies the system fault response when the nominal function defined in SUB-REQ-015 is not achieved within the specified performance threshold. | 2026-03-31T10:14:23.120Z
SUB-POWERSUBSYSTEM-002 | SUB-POWERSUBSYSTEM-009 | derives | SUB-REQ-014 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-014: it specifies the system fault response when the nominal function defined in SUB-REQ-014 is not achieved within the specified performance threshold. | 2026-03-31T10:14:22.343Z
SUB-POWERSUBSYSTEM-004 | SUB-POWERSUBSYSTEM-011 | derives | SUB-REQ-016 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-016: it specifies the system fault response when the nominal function defined in SUB-REQ-016 is not achieved within the specified performance threshold. | 2026-03-31T10:14:21.513Z
SUB-PERCEPTIONSUBSYSTEM-006 | SUB-POWERSUBSYSTEM-013 | derives | SUB-REQ-040 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-040: it specifies the system fault response when the nominal function defined in SUB-REQ-040 is not achieved within the specified performance threshold. | 2026-03-31T10:14:20.740Z
SUB-POWERSUBSYSTEM-007 | SUB-POWERSUBSYSTEM-015 | derives | SUB-REQ-071 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-071: it specifies the system fault response when the nominal function defined in SUB-REQ-071 is not achieved within the specified performance threshold. | 2026-03-31T10:14:19.930Z
SUB-PERCEPTIONSUBSYSTEM-001 | SUB-PERCEPTIONSUBSYSTEM-008 | derives | SUB-REQ-023 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-023: it specifies the system fault response when the nominal function defined in SUB-REQ-023 is not achieved within the specified performance threshold. | 2026-03-31T10:14:19.145Z
SUB-PERCEPTIONSUBSYSTEM-002 | SUB-PERCEPTIONSUBSYSTEM-009 | derives | SUB-REQ-024 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-024: it specifies the system fault response when the nominal function defined in SUB-REQ-024 is not achieved within the specified performance threshold. | 2026-03-31T10:14:18.369Z
SUB-PERCEPTIONSUBSYSTEM-003 | SUB-PERCEPTIONSUBSYSTEM-010 | derives | SUB-REQ-025 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-025: it specifies the system fault response when the nominal function defined in SUB-REQ-025 is not achieved within the specified performance threshold. | 2026-03-31T10:14:17.592Z
SUB-VEHICLEPLATFORM-001 | SUB-VEHICLEPLATFORM-010 | derives | SUB-REQ-027 nominal → failure-mode companion | The failure-mode requirement derives from SUB-REQ-027: it specifies the system fault response when the nominal function defined in SUB-REQ-027 is not achieved within the specified performance threshold. | 2026-03-31T10:14:16.811Z
SUB-POWERSUBSYSTEM-001 | SUB-POWERSUBSYSTEM-008 | derives | BMS disconnect timeout → Safety Monitor Processor emergency stop | SUB-REQ-080 specifies the Safety Monitor Processor fault response when the BMS disconnect function in SUB-REQ-013 fails within the 200ms threshold. The failure-mode requirement exists because the 200ms disconnect is the nominal performance whose violation initiates thermal runaway risk. | 2026-03-31T10:13:59.776Z
STK-REQ-017 | SYS-REQ-021 | derives | STK-REQ-017 outdoor environmental use → SYS-REQ-021 IP54 enclosure requirement | STK-REQ-017 requires safe outdoor operation in ambient temperatures -5°C to 40°C, which implies protection against water and particulate ingress in uncontrolled environments. SYS-REQ-021 derives this into an IP54 sealed Electronics Bay requirement — IP54 is the minimum enclosure standard for outdoor light-use conditions (IEC 60529) and ensures electronics survive rain and humidity exposure consistent with STK-REQ-017's outdoor operating environment. | 2026-03-31T09:40:07.188Z
STK-REQ-002 | SYS-REQ-011 | derives | STK-REQ-002 user safety → SYS-REQ-011 tilt/tip-over protection | STK-REQ-002 mandates user injury protection across all operating modes including tilt/rollover scenarios. SYS-REQ-011 derives from this by specifying the 15° inclinometer threshold that triggers braking before the vehicle reaches tip-over angle (~25°), directly implementing H-005 (tip-over) mitigation required by STK-REQ-002's safety mandate. | 2026-03-31T09:40:06.396Z
VER-BCIPROCESSINGSUBSYSTEM-017 | SYS-REQ-008 | verifies | System-level integration test for cognitive fatigue mode transition | REQ-SEEEGACCESSIBILITYBUGGY-002 directly verifies the end-to-end behaviour specified in SYS-REQ-008: rolling 2-minute accuracy calculation, 70% threshold crossing, mode transition, audible alert generation, and speed limitation. This closes the validation gap where SYS-REQ-008 had only subsystem-level verification. | 2026-03-31T08:35:03.414Z
VER-BCIPROCESSINGSUBSYSTEM-016 | SYS-REQ-007 | verifies | System-level integration test for SNR degradation mode transition | REQ-SEEEGACCESSIBILITYBUGGY-001 directly verifies the end-to-end behaviour specified in SYS-REQ-007: SNR threshold detection, 5-second timing, mode transition, speed clamping, and command simplification. This closes the validation gap where SYS-REQ-007 had only subsystem-level verification. | 2026-03-31T08:35:01.316Z
VER-VEHICLEPLATFORM-012 | IFC-VEHICLEPLATFORM-002 | verifies | VER-REQ-114 verifies IFC-REQ-031 wheel quick-release | VER-REQ-114 tests release force (≤15N), interlock against partial engagement, and engagement confirmation specified in IFC-REQ-031. | 2026-03-31T08:13:26.486Z
VER-HMISUBSYSTEM-014 | IFC-HMISUBSYSTEM-001 | verifies | VER-REQ-113 verifies IFC-REQ-026 MAP to display SPI | VER-REQ-113 tests SPI clock frequency (40MHz), frame rate (≥30fps), and chip-select integrity specified in IFC-REQ-026. | 2026-03-31T08:13:25.668Z
VER-COMMUNICATIONSUBSYSTEM-013 | IFC-COMMUNICATIONSUBSYSTEM-003 | verifies | VER-REQ-112 verifies IFC-REQ-025 UART to BLE module | VER-REQ-112 tests UART baud rate, HCI transfer latency, and hardware flow control behaviour specified in IFC-REQ-025. | 2026-03-31T08:13:24.787Z
VER-VEHICLEPLATFORM-011 | IFC-VEHICLEPLATFORM-001 | verifies | VER-REQ-111 verifies IFC-REQ-024 Electronics Bay chassis fasteners | VER-REQ-111 tests pull-out strength (≥500N/fastener), vibration isolator integrity, and mechanical interlock specified in IFC-REQ-024 through axial pull and shock tests. | 2026-03-31T08:13:23.988Z
VER-COMMUNICATIONSUBSYSTEM-012 | IFC-COMMUNICATIONSUBSYSTEM-002 | verifies | VER-REQ-110 verifies IFC-REQ-006 CMMS REST API interface | VER-REQ-110 tests CMMS REST API data delivery, authentication rejection, and TLS 1.3 enforcement specified in IFC-REQ-006. | 2026-03-31T08:13:13.130Z
VER-COMMUNICATIONSUBSYSTEM-011 | IFC-COMMUNICATIONSUBSYSTEM-001 | verifies | VER-REQ-109 verifies IFC-REQ-004 BLE companion app interface | VER-REQ-109 tests BLE 5.0 companion app connectivity, emergency alert latency, and AES-128 link-layer encryption specified in IFC-REQ-004. | 2026-03-31T08:13:12.300Z
VER-COMMUNICATIONSUBSYSTEM-010 | SUB-COMMUNICATIONSUBSYSTEM-006 | verifies | VER-REQ-108 verifies SUB-REQ-056 comm subsystem diagnostic API | VER-REQ-108 tests the communication subsystem diagnostic API response time and log timestamping specified in SUB-REQ-056. | 2026-03-31T08:13:11.548Z
VER-BCIPROCESSINGSUBSYSTEM-015 | SUB-BCIPROCESSINGSUBSYSTEM-007 | verifies | VER-REQ-107 verifies SUB-REQ-037 USB-C diagnostic port | VER-REQ-107 tests the USB-C service port diagnostic test suite completion time and reporting completeness specified in SUB-REQ-037. | 2026-03-31T08:13:10.706Z
VER-VEHICLEPLATFORM-010 | SUB-VEHICLEPLATFORM-003 | verifies | VER-REQ-106 verifies SUB-REQ-029 seat and occupant restraint | VER-REQ-106 tests the occupant weight/stature range, backrest recline angle, and 4-point harness restraint performance specified in SUB-REQ-029. | 2026-03-31T08:13:09.808Z
VER-VEHICLEPLATFORM-009 | SUB-VEHICLEPLATFORM-008 | verifies | VER-REQ-105 verifies SUB-REQ-075 Electronics Bay IP54 and thermal | VER-REQ-105 tests IP54 ingress protection, operating temperature ceiling, and maintenance cycle durability of the Electronics Bay per SUB-REQ-075. | 2026-03-31T08:12:56.757Z
VER-POWERSUBSYSTEM-010 | SUB-POWERSUBSYSTEM-006 | verifies | VER-REQ-104 verifies SUB-REQ-055 charging parameters | VER-REQ-104 directly tests the inrush current limit, cell temperature ceiling, and charge time specified in SUB-REQ-055. | 2026-03-31T08:12:55.958Z
VER-REQ-103 | SUB-BCIPROCESSINGSUBSYSTEM-013 | verifies | VER-REQ-103 verifies SUB-REQ-079 command authentication | VER-REQ-103 tests the HMAC-SHA256 command authentication specified in SUB-REQ-079 using deliberate signature tampering and nonce replay to confirm cryptographic rejection end-to-end. | 2026-03-31T08:12:55.122Z
VER-PERCEPTIONSUBSYSTEM-010 | IFC-PERCEPTIONSUBSYSTEM-001 | verifies | Verification coverage for IFC-REQ-021 (depth sensor I2C bus, SIL-2) | VER-REQ-102 tests I2C enumeration, latency and bus integrity for all three forward depth sensors | 2026-03-31T07:10:50.088Z
VER-DRIVESUBSYSTEM-013 | SUB-DRIVESUBSYSTEM-010 | verifies | Verification coverage for SUB-REQ-077 (drive subsystem IEC 60601-1 EMC compliance) | VER-REQ-101 requires accredited lab EMC testing per IEC 60601-1-2 for the drive subsystem | 2026-03-31T07:09:55.056Z
VER-POWERSUBSYSTEM-009 | SUB-POWERSUBSYSTEM-005 | verifies | Verification coverage for SUB-REQ-036 (charge controller dual-mains AC input) | VER-REQ-100 confirms charge acceptance at both 230V/50Hz and 120V/60Hz facility supplies | 2026-03-31T07:09:41.321Z
VER-PERCEPTIONSUBSYSTEM-009 | IFC-PERCEPTIONSUBSYSTEM-002 | verifies | Verification coverage for IFC-REQ-022 (lateral proximity sensor GPIO interface) | VER-REQ-099 tests detection boundaries and timing at both margins for the SIL-2 lateral obstacle GPIO | 2026-03-31T07:09:27.605Z
VER-POWERSUBSYSTEM-008 | IFC-POWERSUBSYSTEM-001 | verifies | Verification coverage for IFC-REQ-002 (charging dock galvanic isolation) | VER-REQ-098 tests isolation resistance and auto-termination for the safety-critical charging interface | 2026-03-31T07:09:15.774Z
VER-COMMUNICATIONSUBSYSTEM-009 | IFC-REQ-005 | verifies | Verification coverage for IFC-REQ-005 (USB-C service port operations) | VER-REQ-097 tests all four maintenance operations on the USB-C service interface | 2026-03-31T07:09:02.511Z
VER-VEHICLEPLATFORM-008 | SUB-VEHICLEPLATFORM-006 | verifies | Verification coverage for SUB-REQ-072 (maintenance diagnostic USB-C) | VER-REQ-096 tests the diagnostic API and port enumeration required by the Weekly Maintenance scenario | 2026-03-31T07:08:49.938Z
VER-PERCEPTIONSUBSYSTEM-008 | SUB-PERCEPTIONSUBSYSTEM-006 | verifies | Verification coverage for SUB-REQ-040 (inclinometer SIL-2 power supply tolerance) | VER-REQ-095 tests power rail tolerance and ripple resilience for SIL-2 tilt sensor | 2026-03-31T07:08:38.547Z
VER-VEHICLEPLATFORM-007 | SUB-VEHICLEPLATFORM-007 | verifies | Verification coverage for SUB-REQ-074 (chassis material and structural integrity, SIL-2) | VER-REQ-094 provides structural proof test and material inspection for SIL-2 chassis frame requirement | 2026-03-31T07:08:24.307Z
VER-HMISUBSYSTEM-013 | SUB-HMISUBSYSTEM-007 | verifies | Verification coverage for SUB-REQ-053 (carer joystick routing latency) | VER-REQ-093 tests the joystick-to-CAN latency required by SUB-REQ-053 | 2026-03-31T07:08:09.823Z
VER-HMISUBSYSTEM-012 | SUB-HMISUBSYSTEM-006 | verifies | Verification coverage for SUB-REQ-052 (carer override assertion timing) | VER-REQ-092 measures the CARER_OVERRIDE assertion timing required by SUB-REQ-052 | 2026-03-31T07:07:57.409Z
VER-BCIPROCESSINGSUBSYSTEM-014 | IFC-BCIPROCESSINGSUBSYSTEM-004 | verifies | Verification coverage for IFC-REQ-013 (SIL-3 feature vector interface) | VER-REQ-091 provides the 1000-iteration latency and integrity test for the SIL-3 BCI classifier input interface | 2026-03-31T07:07:46.435Z
VER-POWERSUBSYSTEM-007 | IFC-POWERSUBSYSTEM-003 | verifies | Verification coverage for IFC-REQ-017 | VER-REQ-090 provides the verification procedure for the interface specification in IFC-REQ-017 | 2026-03-31T06:39:54.512Z
VER-BCIPROCESSINGSUBSYSTEM-013 | IFC-BCIPROCESSINGSUBSYSTEM-005 | verifies | Verification coverage for IFC-REQ-014 | VER-REQ-089 provides the verification procedure for the interface specification in IFC-REQ-014 | 2026-03-31T06:39:53.681Z
VER-BCIPROCESSINGSUBSYSTEM-012 | IFC-BCIPROCESSINGSUBSYSTEM-003 | verifies | Verification coverage for IFC-REQ-012 | VER-REQ-088 provides the verification procedure for the interface specification in IFC-REQ-012 | 2026-03-31T06:39:52.829Z
VER-SAFETYSUBSYSTEM-023 | IFC-SAFETYSUBSYSTEM-005 | verifies | Verification coverage for IFC-REQ-010 | VER-REQ-087 provides the verification procedure for the interface specification in IFC-REQ-010 | 2026-03-31T06:39:51.959Z
VER-BCIPROCESSINGSUBSYSTEM-011 | SUB-BCIPROCESSINGSUBSYSTEM-012 | verifies | Verification coverage for SUB-REQ-066 | VER-REQ-086 provides the verification procedure for the compliance or safety requirement in SUB-REQ-066 | 2026-03-31T06:38:51.698Z
VER-DRIVESUBSYSTEM-012 | SUB-DRIVESUBSYSTEM-008 | verifies | Verification coverage for SUB-REQ-065 | VER-REQ-085 provides the verification procedure for the compliance or safety requirement in SUB-REQ-065 | 2026-03-31T06:38:50.865Z
VER-DRIVESUBSYSTEM-011 | SUB-DRIVESUBSYSTEM-007 | verifies | Verification coverage for SUB-REQ-064 | VER-REQ-084 provides the verification procedure for the compliance or safety requirement in SUB-REQ-064 | 2026-03-31T06:38:50.014Z
VER-SAFETYSUBSYSTEM-022 | SUB-SAFETYSUBSYSTEM-011 | verifies | Verification coverage for SUB-REQ-059 | VER-REQ-083 provides the verification procedure for the compliance or safety requirement in SUB-REQ-059 | 2026-03-31T06:38:49.197Z
VER-VEHICLEPLATFORM-006 | SUB-VEHICLEPLATFORM-005 | verifies | Verification coverage for SUB-REQ-050 | VER-REQ-082 provides the verification procedure for the compliance or safety requirement in SUB-REQ-050 | 2026-03-31T06:38:48.325Z
VER-COMMUNICATIONSUBSYSTEM-008 | SUB-COMMUNICATIONSUBSYSTEM-005 | verifies | Verification coverage for SUB-REQ-046 | VER-REQ-081 provides the verification procedure for the compliance or safety requirement in SUB-REQ-046 | 2026-03-31T06:38:47.447Z
VER-COMMUNICATIONSUBSYSTEM-007 | SUB-COMMUNICATIONSUBSYSTEM-004 | verifies | Verification coverage for SUB-REQ-045 | VER-REQ-080 provides the verification procedure for the compliance or safety requirement in SUB-REQ-045 | 2026-03-31T06:38:46.620Z
VER-BCIPROCESSINGSUBSYSTEM-010 | SUB-BCIPROCESSINGSUBSYSTEM-011 | verifies | Verification coverage for SUB-REQ-044 | VER-REQ-079 provides the verification procedure for the compliance or safety requirement in SUB-REQ-044 | 2026-03-31T06:38:45.834Z
VER-BCIPROCESSINGSUBSYSTEM-009 | SUB-BCIPROCESSINGSUBSYSTEM-009 | verifies | Verification coverage for SUB-REQ-041 | VER-REQ-078 provides the verification procedure for the performance or safety requirement in SUB-REQ-041 | 2026-03-31T06:38:40.980Z
VER-BCIPROCESSINGSUBSYSTEM-008 | SUB-BCIPROCESSINGSUBSYSTEM-008 | verifies | Verification coverage for SUB-REQ-039 | VER-REQ-077 provides the verification procedure for the performance or safety requirement in SUB-REQ-039 | 2026-03-31T06:38:40.162Z
VER-HMISUBSYSTEM-011 | SUB-HMISUBSYSTEM-002 | verifies | Verification coverage for SUB-REQ-033 | VER-REQ-076 provides the verification procedure for the performance or safety requirement in SUB-REQ-033 | 2026-03-31T06:38:39.322Z
VER-COMMUNICATIONSUBSYSTEM-006 | SUB-COMMUNICATIONSUBSYSTEM-001 | verifies | Verification coverage for SUB-REQ-030 | VER-REQ-075 provides the verification procedure for the performance or safety requirement in SUB-REQ-030 | 2026-03-31T06:38:38.538Z
VER-VEHICLEPLATFORM-005 | SUB-VEHICLEPLATFORM-002 | verifies | Verification coverage for SUB-REQ-028 | VER-REQ-074 provides the verification procedure for the performance or safety requirement in SUB-REQ-028 | 2026-03-31T06:38:37.700Z
VER-PERCEPTIONSUBSYSTEM-007 | SUB-PERCEPTIONSUBSYSTEM-003 | verifies | Verification coverage for SUB-REQ-025 | VER-REQ-073 provides the verification procedure for the performance or safety requirement in SUB-REQ-025 | 2026-03-31T06:38:36.947Z
VER-PERCEPTIONSUBSYSTEM-006 | SUB-PERCEPTIONSUBSYSTEM-002 | verifies | Verification coverage for SUB-REQ-024 | VER-REQ-072 provides the verification procedure for the performance or safety requirement in SUB-REQ-024 | 2026-03-31T06:38:36.181Z
VER-DRIVESUBSYSTEM-010 | SUB-DRIVESUBSYSTEM-005 | verifies | Verification coverage for SUB-REQ-022 | VER-REQ-071 provides the verification procedure for the performance or safety requirement in SUB-REQ-022 | 2026-03-31T06:38:35.370Z
VER-DRIVESUBSYSTEM-009 | SUB-DRIVESUBSYSTEM-004 | verifies | Verification coverage for SUB-REQ-021 | VER-REQ-070 provides the verification procedure for the quantified performance requirement in SUB-REQ-021 | 2026-03-31T06:38:30.245Z
VER-DRIVESUBSYSTEM-008 | SUB-DRIVESUBSYSTEM-002 | verifies | Verification coverage for SUB-REQ-019 | VER-REQ-069 provides the verification procedure for the quantified performance requirement in SUB-REQ-019 | 2026-03-31T06:38:29.444Z
VER-DRIVESUBSYSTEM-007 | SUB-DRIVESUBSYSTEM-001 | verifies | Verification coverage for SUB-REQ-018 | VER-REQ-068 provides the verification procedure for the quantified performance requirement in SUB-REQ-018 | 2026-03-31T06:38:28.643Z
VER-BCIPROCESSINGSUBSYSTEM-007 | SUB-BCIPROCESSINGSUBSYSTEM-006 | verifies | Verification coverage for SUB-REQ-017 | VER-REQ-067 provides the verification procedure for the quantified performance requirement in SUB-REQ-017 | 2026-03-31T06:38:27.856Z
VER-POWERSUBSYSTEM-006 | SUB-POWERSUBSYSTEM-004 | verifies | Verification coverage for SUB-REQ-016 | VER-REQ-066 provides the verification procedure for the quantified performance requirement in SUB-REQ-016 | 2026-03-31T06:38:27.058Z
VER-POWERSUBSYSTEM-005 | SUB-POWERSUBSYSTEM-003 | verifies | Verification coverage for SUB-REQ-015 | VER-REQ-065 provides the verification procedure for the quantified performance requirement in SUB-REQ-015 | 2026-03-31T06:38:26.259Z
VER-POWERSUBSYSTEM-004 | SUB-POWERSUBSYSTEM-002 | verifies | Verification coverage for SUB-REQ-014 | VER-REQ-064 provides the verification procedure for the quantified performance requirement in SUB-REQ-014 | 2026-03-31T06:38:25.430Z
VER-BCIPROCESSINGSUBSYSTEM-006 | SUB-BCIPROCESSINGSUBSYSTEM-005 | verifies | Verification coverage for SUB-REQ-012 | VER-REQ-063 provides the verification procedure for the quantified performance requirement in SUB-REQ-012 | 2026-03-31T06:38:24.597Z
SYS-REQ-019 | SUB-BCIPROCESSINGSUBSYSTEM-013 | derives | Command authentication derives from system-level access control requirement | SYS-REQ-019 mandates role-based access control for interfaces; SUB-REQ-079 decomposes this to the MAP level, specifically protecting BCI navigation command integrity from external interface injection | 2026-03-31T05:39:35.311Z
VER-VEHICLEPLATFORM-004 | SUB-VEHICLEPLATFORM-009 | verifies | Chassis structural compliance verification | VER-REQ-062 provides ISO 7176-8 static and impact test procedure for SUB-REQ-078 chassis structural compliance | 2026-03-31T05:38:52.820Z
SYS-REQ-010 | SUB-VEHICLEPLATFORM-009 | derives | Chassis structural compliance derives from system mass and dimensional constraint | SYS-REQ-010 specifies vehicle mass constraints; SUB-REQ-078 decomposes the structural requirement ensuring the chassis can bear those loads with a safety margin per ISO 7176-8 | 2026-03-31T05:38:45.523Z
SYS-REQ-013 | SUB-DRIVESUBSYSTEM-010 | derives | Drive subsystem EMC compliance derives from system-level medical device EMC requirement | SYS-REQ-013 requires IEC 60601-1-2 EMC compliance at system level; SUB-REQ-077 decomposes this to the Drive Subsystem components where 48V switching PWM is the primary EMC emission source | 2026-03-31T05:38:22.900Z
VER-SAFETYSUBSYSTEM-021 | SYS-REQ-023 | verifies | System-level MAP failsafe verification | VER-REQ-061 provides system-level integration test for SYS-REQ-023 MAP watchdog failsafe | 2026-03-31T05:37:41.059Z
SYS-REQ-023 | SUB-SAFETYSUBSYSTEM-013 | derives | SMP MAP-failure detection derives from SYS redundancy requirement | SYS-REQ-023 requires SMP to detect and respond to MAP failure; SUB-REQ-061 implements this with watchdog timeout, CRC fault detection, and SAFE_STOP assertion | 2026-03-31T05:37:30.923Z
STK-REQ-002 | SYS-REQ-023 | derives | MAP watchdog failover derives from user safety stakeholder requirement | STK-REQ-002 mandates user safety via reliable emergency stop; SYS-REQ-023 addresses the specific failure mode of MAP fault leading to loss of emergency stop capability without independent detection | 2026-03-31T05:37:30.120Z
VER-DRIVESUBSYSTEM-006 | SUB-DRIVESUBSYSTEM-009 | verifies | Override mode performance verification | VER-REQ-060 provides direct test procedure for SUB-REQ-076 override mode performance criteria | 2026-03-31T05:35:14.937Z
SYS-REQ-009 | SUB-DRIVESUBSYSTEM-009 | derives | Override mode performance derives from carer handover requirement | SYS-REQ-009 requires 100ms handover to joystick on override activation; SUB-REQ-076 decomposes the quantified performance floor for that override mode including sustained speed, obstacle detection response, and endurance | 2026-03-31T05:35:07.384Z
SYS-REQ-021 | SUB-VEHICLEPLATFORM-008 | derives | SYS-REQ-021 sealed Electronics Bay requirement → Electronics Bay material and environmental requirement | SUB-REQ-075 decomposes the Electronics Bay aspect of SYS-REQ-021, specifying IP54 enclosure, thermal management bounds, and vibration resistance to satisfy the sealed bay requirement of the integrated chassis. | 2026-03-31T05:11:26.381Z
SYS-REQ-021 | SUB-VEHICLEPLATFORM-007 | derives | SYS-REQ-021 integrated chassis housing → Chassis Frame material and structural requirement | SUB-REQ-074 decomposes the physical chassis aspect of SYS-REQ-021, specifying the material, wall thickness, load capacity, and corrosion resistance that the single integrated chassis enclosure must provide. | 2026-03-31T05:11:25.540Z
SYS-REQ-007 | SUB-BCIPROCESSINGSUBSYSTEM-012 | derives | SYS-REQ-007 BCI SNR degradation threshold → Artifact Rejection Engine CPU/latency resource bounds | SUB-REQ-066 derives from SYS-REQ-007 because the SNR floor (signal quality degradation trigger) is maintained by the Artifact Rejection Engine. Its CPU and latency constraints are driven by the need to process artifacts in real-time without delaying the classification pipeline that detects SNR degradation. | 2026-03-31T05:11:17.090Z
VER-HMISUBSYSTEM-010 | SUB-HMISUBSYSTEM-011 | verifies | Verification of amber LED and display text notification in degraded mode with timing | Direct observation test with timing measurement confirms the visual HMI requirements are met. Text string match ensures care communication is unambiguous. | 2026-03-31T05:10:48.952Z
VER-POWERSUBSYSTEM-003 | SUB-POWERSUBSYSTEM-007 | verifies | Verification of CC-CV charging profile and 4-hour completion on both mains standards | Full charge cycle test with measured timing and SOC confirmation is the minimum viable verification for the charging specification. Dual supply test verifies international deployment capability. | 2026-03-31T05:10:38.726Z
VER-SAFETYSUBSYSTEM-020 | SUB-SAFETYSUBSYSTEM-015 | verifies | Verification of inclinometer dual-threshold detection and brake response timing | Tilt platform test provides controlled, repeatable tilt angle increments to validate both the pre-warning threshold and the hard-stop threshold with measured timing precision. | 2026-03-31T05:10:21.866Z
VER-HMISUBSYSTEM-009 | SUB-HMISUBSYSTEM-010 | verifies | Verification of joystick override handover timing and BCI command interlock | 100ms timing verification and BCI command rejection during override confirm the safety interlock is working. The 3 in-flight BCI command trials are the critical edge case that validates the authority transfer is robust under actual operating conditions. | 2026-03-31T05:10:04.482Z
VER-HMISUBSYSTEM-008 | SUB-HMISUBSYSTEM-009 | verifies | Verification of Audio Alert Module sound level and response time on BCI accuracy threshold | Sound level meter measurement at 1m provides direct verification of 85 dBSPL specification. Response time measurement from threshold crossing to tone onset verifies the 200ms requirement. | 2026-03-31T05:09:46.663Z
VER-COMMUNICATIONSUBSYSTEM-005 | SUB-COMMUNICATIONSUBSYSTEM-007 | verifies | Verification of BLE emergency alert transmission timing and content | Pass/fail criteria: packet received within 500ms with correct fields on 10 consecutive trials. This is sufficient to demonstrate that the Communication Controller meets the alert transmission requirement under representative operating conditions. | 2026-03-31T05:09:31.779Z
SYS-REQ-022 | SUB-HMISUBSYSTEM-011 | derives | SYS-REQ-022 amber LED and text display in degraded mode → HMI Status LED Array and Display Unit requirement | SUB-REQ-073 decomposes SYS-REQ-022's visual status requirements into specific LED pulse rate, display text content, and 100ms latency constraints for the HMI Subsystem components. | 2026-03-31T05:09:01.181Z
SYS-REQ-015 | SUB-VEHICLEPLATFORM-006 | derives | SYS-REQ-015 automated diagnostic via USB-C → Vehicle Platform Electronics Bay service port requirement | SUB-REQ-072 decomposes SYS-REQ-015 by specifying the USB-C connector standard, each diagnostic metric's acceptance threshold, and the 20-minute completion constraint for the Electronics Bay service port. | 2026-03-31T05:08:59.982Z
SYS-REQ-014 | SUB-POWERSUBSYSTEM-007 | derives | SYS-REQ-014 facility charging dock 4-hour charge → Power Subsystem Charge Controller requirement | SUB-REQ-071 decomposes SYS-REQ-014 by specifying the CC-CV charging profile, dual-voltage acceptance, and charge termination behaviour of the Charge Controller that achieves the 4-hour 20%-100% SOC system requirement. | 2026-03-31T05:08:59.165Z
SYS-REQ-011 | SUB-SAFETYSUBSYSTEM-015 | derives | SYS-REQ-011 inclinometer tilt safety → Safety Subsystem Inclinometer Tilt Sensor requirement | SUB-REQ-070 decomposes SYS-REQ-011's 15-degree tilt braking requirement into the specific two-threshold detection scheme (12° pre-warning, 15° hard stop) and 50ms response time requirement for the Inclinometer Tilt Sensor Unit. | 2026-03-31T05:08:48.455Z
SYS-REQ-009 | SUB-HMISUBSYSTEM-010 | derives | SYS-REQ-009 care attendant joystick override → HMI Subsystem handover requirement | SUB-REQ-069 decomposes SYS-REQ-009's 100ms joystick authority transfer requirement into the HMI Subsystem's specific signal propagation, BCI disable, and Drive Subsystem confirmation steps. | 2026-03-31T05:08:47.544Z
SYS-REQ-008 | SUB-HMISUBSYSTEM-009 | derives | SYS-REQ-008 degraded mode care attendant alert → HMI Audio Alert Module requirement | SUB-REQ-068 decomposes SYS-REQ-008's requirement for audible alert to care attendant, specifying the Audio Alert Module's dB level, frequency, and timing constraints to achieve the system-level notification requirement. | 2026-03-31T05:08:46.640Z
SYS-REQ-005 | SUB-COMMUNICATIONSUBSYSTEM-007 | derives | SYS-REQ-005 seizure emergency BLE alert → Communication Subsystem BLE transmission requirement | SUB-REQ-067 exists specifically because SYS-REQ-005 requires BLE emergency alert to the facility emergency system; the subsystem requirement decomposes the specific transmission protocol, timing, and content of that alert. | 2026-03-31T05:08:35.304Z
VER-VEHICLEPLATFORM-003 | SYS-REQ-021 | verifies | IEC 60529 IP54 test verifies physical housing protection | Standardised IP54 test procedure provides objective evidence for the environmental protection claim | 2026-03-31T04:38:21.088Z
VER-SAFETYSUBSYSTEM-019 | SUB-SAFETYSUBSYSTEM-014 | verifies | SIL-3 certification document review verifies functional safety compliance | Third-party IEC 61508 assessment is required for SIL-3 claims; document inspection is the only applicable verification method | 2026-03-31T04:38:07.512Z
VER-SAFETYSUBSYSTEM-018 | SUB-SAFETYSUBSYSTEM-013 | verifies | MAP watchdog fault injection test verifies SMP failover independence | Fault injection is the only reliable way to verify that the SMP failover path operates without MAP firmware participation | 2026-03-31T04:37:51.881Z
VER-SAFETYSUBSYSTEM-017 | SUB-SAFETYSUBSYSTEM-012 | verifies | Per-channel relay drop-out test verifies dual-channel E-stop independence | The safety claim requires single-channel sufficiency; only independent channel tests can confirm this. | 2026-03-31T04:37:40.829Z
SYS-REQ-022 | SUB-DRIVESUBSYSTEM-008 | derives | Drive subsystem 3 km/h degraded speed cap derives from system degraded mode performance requirement | SYS-REQ-022 mandates 3 km/h maximum in Degraded mode; this SUB requirement specifies the mechanism (PWM duty cycle cap) and its independence from input source | 2026-03-31T04:37:27.017Z
STK-REQ-005 | SYS-REQ-022 | derives | Amber LED alert in Degraded mode derives from stakeholder alert requirement | STK-REQ-005 requires distinct audible and visual signals on Degraded mode entry; this system requirement specifies the amber LED and display notification as the distinct visual signal | 2026-03-31T04:37:15.648Z
STK-REQ-004 | SYS-REQ-022 | derives | Degraded mode 3 km/h speed limit derives from stakeholder care attendant override requirement | STK-REQ-004 requires manual override with obstacle detection active; this system requirement adds the measurable speed constraint and alert requirement missing from the STK | 2026-03-31T04:37:14.856Z
SYS-REQ-002 | SUB-DRIVESUBSYSTEM-007 | derives | MCU hardware overcurrent protection derives from safety trigger emergency stop requirement | SYS-REQ-002 requires emergency stop on any safety trigger; this requirement provides a hardware-layer safety net for motor overcurrent, independent of software fault detection | 2026-03-31T04:36:54.673Z
VER-HMISUBSYSTEM-007 | SUB-HMISUBSYSTEM-008 | verifies | Verification of 100ms joystick authority transfer timing | Pass/fail criterion of 100ms max with oscilloscope measurement provides traceable evidence that the authority transfer budget is met under operational conditions | 2026-03-31T04:36:12.743Z
SYS-REQ-009 | SUB-HMISUBSYSTEM-008 | derives | 100ms joystick authority transfer timing derives from care attendant override requirement | SYS-REQ-009 specifies within 100ms; SUB-REQ-062 decomposes this budget into measurable hardware/software timing stages at the HMI subsystem boundary | 2026-03-31T04:35:57.665Z
SYS-REQ-002 | SUB-SAFETYSUBSYSTEM-013 | derives | MAP failover to Safety Monitor Processor derives from safety trigger emergency stop | SYS-REQ-002 requires emergency stop on system fault; SUB-REQ-061 ensures MAP processor failure is treated as a system fault requiring immediate failover to the SMP | 2026-03-31T04:35:44.170Z
SYS-REQ-002 | SUB-SAFETYSUBSYSTEM-012 | derives | Dual-channel E-stop circuit derives from safety trigger emergency stop requirement | SYS-REQ-002 mandates emergency stop on any safety trigger; this requirement ensures no single-point-of-failure in the physical circuit, which is required for SIL-3 compliance | 2026-03-31T04:35:15.378Z
SYS-REQ-005 | SUB-COMMUNICATIONSUBSYSTEM-003 | derives | Communication controller BLE alert derives from seizure-triggered emergency transmission | SYS-REQ-005 requires BLE emergency alert on seizure; SUB-REQ-038 implements this at the Communication Controller level | 2026-03-31T04:34:20.612Z
VER-PERCEPTIONSUBSYSTEM-005 | SUB-PERCEPTIONSUBSYSTEM-007 | verifies | Physical tilt test verifying inclinometer publishing rate and content | Physical tilt on test fixture with bus monitoring validates both rate and content; 30 s duration confirms sustained operation above threshold | 2026-03-31T04:10:13.227Z
VER-DRIVESUBSYSTEM-005 | SUB-DRIVESUBSYSTEM-006 | verifies | CAN failure injection test verifying MCU heartbeat monitoring | Hardware-in-the-loop CAN failure injection validates both timing and state transition without degrading production hardware | 2026-03-31T04:10:12.304Z
VER-SAFETYSUBSYSTEM-016 | SUB-SAFETYSUBSYSTEM-010 | verifies | Watchdog injection test verifying MAP failover chain | Direct watchdog injection is the only reliable method to verify timing without simulation; NV log check confirms fault record persistence | 2026-03-31T04:10:11.516Z
VER-SAFETYSUBSYSTEM-015 | SUB-SAFETYSUBSYSTEM-009 | verifies | Integration test verifying facility E-stop response timing | Timed test with live RF link validates the full hardware path; pass criterion of 150 ms over 20 trials provides statistical confidence in SIL-3 response timing | 2026-03-31T04:10:10.714Z
SYS-REQ-010 | ARC-REQ-007 | derives | Vehicle Platform four-component decomposition driven by physical footprint constraint | SYS-REQ-010 constrains platform to 700x1200 mm; ARC-REQ-007 records the structural decomposition that meets the footprint while integrating occupant support and all electronics | 2026-03-31T04:09:31.428Z
SYS-REQ-003 | ARC-REQ-006 | derives | Differential-drive BLDC architecture driven by speed and control requirements | SYS-REQ-003 constrains speed to 6 km/h; ARC-REQ-006 records the differential-drive decision which provides steering without steering actuators, simplifying speed control | 2026-03-31T04:09:30.634Z
SYS-REQ-012 | ARC-REQ-005 | derives | BMS co-location architecture driven by battery thermal safety requirement | SYS-REQ-012 requires power cutoff on thermal fault; ARC-REQ-005 records BMS co-location with safety arbitration to minimise latency in thermal protection chain | 2026-03-31T04:09:29.844Z
SYS-REQ-017 | ARC-REQ-004 | derives | Communication Subsystem separation driven by BLE notification requirement | SYS-REQ-017 requires BLE notification on emergency; ARC-REQ-004 records the Communication-HMI separation to isolate wireless reliability from user feedback display | 2026-03-31T04:09:17.077Z
SYS-REQ-006 | ARC-REQ-003 | derives | Perception-Drive separation driven by obstacle detection requirement | SYS-REQ-006 requires 2m obstacle detection for halt; ARC-REQ-003 records the Perception separation for failure independence from Drive | 2026-03-31T04:09:16.231Z
SYS-REQ-002 | ARC-REQ-002 | derives | Independent safety channel architecture driven by halt requirement | SYS-REQ-002 mandates halt on all safety triggers; ARC-REQ-002 records the dedicated safety processor decision to avoid common-cause failure | 2026-03-31T04:09:15.397Z
SYS-REQ-001 | ARC-REQ-001 | derives | BCI combined processor architecture motivated by classification accuracy | SYS-REQ-001 sets 80% classification accuracy; ARC-REQ-001 records the co-location decision to meet latency constraints for accurate classification | 2026-03-31T04:09:14.313Z
SYS-REQ-013 | SUB-SAFETYSUBSYSTEM-011 | derives | Safety Subsystem SMP must comply with IEC 61508 SIL 3 | SYS-REQ-013 establishes IEC 60601-1-2 as the top-level compliance standard; SUB-REQ-059 extends this to IEC 61508-3 SIL 3 for the Safety Monitor Processor, which is the most critical safety-essential software component | 2026-03-31T04:07:58.834Z
SYS-REQ-002 | SUB-DRIVESUBSYSTEM-006 | derives | Drive Subsystem implements MCU heartbeat monitoring as fault detection | SYS-REQ-002 requires response to system faults; SUB-REQ-058 decomposes MCU heartbeat loss as the Drive Subsystem fault detection mechanism that feeds the safety response chain | 2026-03-31T04:07:57.952Z
SYS-REQ-002 | SUB-SAFETYSUBSYSTEM-010 | derives | Safety Subsystem implements processor watchdog as a system fault trigger | SYS-REQ-002 lists system fault among safety triggers; SUB-REQ-057 decomposes the MAP watchdog fault detection and response as the Safety Subsystem implementation of that trigger | 2026-03-31T04:07:57.204Z
SYS-REQ-015 | SUB-COMMUNICATIONSUBSYSTEM-006 | derives | Communication Subsystem hosts USB-C diagnostic service port | SYS-REQ-015 specifies USB-C diagnostic access; SUB-REQ-056 decomposes this to the Communication Subsystem as the interface host | 2026-03-31T04:06:53.100Z
SYS-REQ-014 | SUB-POWERSUBSYSTEM-006 | derives | Power Subsystem implements facility charging with thermal and current management | SYS-REQ-014 sets 4-hour charge requirement; SUB-REQ-055 decomposes this to the Power Subsystem charger with specific current and thermal constraints | 2026-03-31T04:06:52.277Z
SYS-REQ-011 | SUB-PERCEPTIONSUBSYSTEM-007 | derives | Perception Subsystem implements inclinometer data publishing to Safety Monitor | SYS-REQ-011 triggers halt on tilt; SUB-REQ-054 decomposes the Perception Subsystem requirement to publish tilt data at rate needed for Safety Subsystem response | 2026-03-31T04:06:51.396Z
SYS-REQ-009 | SUB-HMISUBSYSTEM-007 | derives | HMI Subsystem implements joystick command routing during carer override | SYS-REQ-009 specifies joystick control transfer; SUB-REQ-053 specifies the latency and rate for joystick command routing as the second part of the override chain | 2026-03-31T04:06:42.050Z
SYS-REQ-009 | SUB-HMISUBSYSTEM-006 | derives | HMI Subsystem implements hardwired carer override assertion | SYS-REQ-009 mandates care attendant steering transfer; SUB-REQ-052 decomposes the HMI hardware path for the override signal | 2026-03-31T04:06:41.207Z
SYS-REQ-005 | SUB-SAFETYSUBSYSTEM-009 | derives | Safety Subsystem implements facility halt command integration | SYS-REQ-005 requires integration with facility emergency systems; SUB-REQ-051 decomposes this to the Safety Subsystem as the hardware actuator for motor de-energisation | 2026-03-31T04:06:40.388Z
SYS-REQ-002 | SUB-BCIPROCESSINGSUBSYSTEM-004 | derives | Safety halt inheritance chain for BCI SNR hard-stop requirement | SYS-REQ-002 requires immediate halt on complete BCI signal loss. SUB-REQ-011 implements the Command Arbitration hard-stop response to BCI SNR drop below threshold (a precursor to complete signal loss). The sil-3 tag on SUB-REQ-011 is inherited from SYS-REQ-002, not from SYS-REQ-007 (sil-1, which only mandates mode transition). The 3-second threshold in SUB-REQ-011 is intentionally shorter than the 5-second threshold in SYS-REQ-007, providing a pre-emptive hard stop before mode transition logic activates. | 2026-03-31T03:39:30.118Z
VER-HMISUBSYSTEM-006 | IFC-HMISUBSYSTEM-003 | verifies | MAP to LED Array SPI interface verification | VER-REQ-044 defines the SPI frame latency and protocol structure test that verifies the 20ms update latency and 6-element addressability in IFC-REQ-030. | 2026-03-31T03:36:44.986Z
VER-HMISUBSYSTEM-005 | SUB-HMISUBSYSTEM-005 | verifies | Status LED Array visibility verification | VER-REQ-043 defines the photometric angular sweep test that verifies the 120-degree arc visibility and 10 mcd intensity threshold in SUB-REQ-048. | 2026-03-31T03:36:44.112Z
VER-REQ-042 | IFC-BCIPROCESSINGSUBSYSTEM-002 | verifies | EEG Acquisition to Artifact Rejection interface verification | VER-REQ-042 defines the SPI bus measurement test that verifies the sample rate, latency, and loss rate criteria in IFC-REQ-011. | 2026-03-31T03:36:43.256Z
VER-REQ-041 | SUB-BCIPROCESSINGSUBSYSTEM-002 | verifies | BCI Classifier accuracy verification | VER-REQ-041 defines the multi-subject session test that verifies the 75% rolling-window classification accuracy threshold in SUB-REQ-009. | 2026-03-31T03:36:42.438Z
VER-REQ-040 | SUB-BCIPROCESSINGSUBSYSTEM-001 | verifies | Artifact rejection engine verification | VER-REQ-040 defines the annotated EEG epoch injection test that directly verifies the rejection rate and false rejection rate criteria in SUB-REQ-008. | 2026-03-31T03:36:41.623Z
SYS-REQ-013 | SUB-VEHICLEPLATFORM-005 | derives | Electronics Bay IP54 requirement derives from medical device EMC/environmental compliance | SYS-REQ-013 requires IEC 60601-1-2 EMC compliance and general medical device environmental standards. SUB-REQ-050 implements the IP54 environmental protection at the Electronics Bay level, which is required by IEC 60601-1 clause 11.6.5 for mobile medical device enclosures. | 2026-03-31T03:33:55.441Z
SYS-REQ-008 | SUB-HMISUBSYSTEM-005 | derives | LED visibility requirement derives from care attendant notification requirement | SYS-REQ-008 requires visual state notification to care attendant. SUB-REQ-048 defines the LED visibility requirements that ensure the status indication is readable by an approaching attendant at normal approach angles. | 2026-03-31T03:33:54.620Z
SYS-REQ-015 | IFC-VEHICLEPLATFORM-002 | derives | Wheel-chassis mechanical interface derives from diagnostic/maintenance system requirement | SYS-REQ-015 requires maintenance tasks accessible without specialist tools. IFC-REQ-031 defines the mechanical interface between Wheel Assembly and Chassis Frame with the tool-free quick-release specification, directly implementing the wheel replacement maintainability requirement. | 2026-03-31T03:33:53.814Z
SYS-REQ-008 | IFC-HMISUBSYSTEM-003 | derives | LED Array interface derives from care attendant alert requirement | SYS-REQ-008 requires audible/visual alert to care attendant on BCI accuracy degradation. IFC-REQ-030 defines the MAP→LED Array interface that communicates vehicle state, directly implementing the visual alert component of this requirement. | 2026-03-31T03:33:52.975Z
SYS-REQ-017 | SUB-COMMUNICATIONSUBSYSTEM-005 | derives | Telemetry rate requirement derives from emergency event notification system requirement | SYS-REQ-017 requires emergency event transmission — SUB-REQ-046 implements the baseline telemetry rate that makes location and state data available at emergency event time. Without regular telemetry, the system has no current-position data to transmit at the emergency event. | 2026-03-31T03:09:40.915Z
SYS-REQ-019 | SUB-COMMUNICATIONSUBSYSTEM-004 | derives | Cellular modem mTLS derives from system TLS 1.3 encryption requirement | SYS-REQ-019 mandates TLS 1.3 for all data in transit. SUB-REQ-045 implements this specifically for the Cellular Modem's LTE data path with mutual authentication. | 2026-03-31T03:09:40.123Z
VER-SAFETYSUBSYSTEM-014 | SYS-REQ-020 | verifies | ISO 10993-1 biocompatibility documentation inspection | VER-REQ-039 verifies SYS-REQ-020 by inspection of the EEG headset manufacturer biocompatibility evaluation per ISO 10993-1. | 2026-03-31T03:09:39.218Z
VER-VEHICLEPLATFORM-002 | SUB-VEHICLEPLATFORM-004 | verifies | Turning radius doorway navigation verification | VER-REQ-038 performs the ISO 7176-3 turning test and doorway simulation to verify the 600mm radius and 150mm clearance margins. | 2026-03-31T03:09:09.956Z
VER-HMISUBSYSTEM-004 | SUB-HMISUBSYSTEM-004 | verifies | Display content and update rate verification | VER-REQ-037 tests the 2Hz update rate and 300 cd/m2 brightness mandated by SUB-REQ-047. | 2026-03-31T03:09:09.184Z
VER-HMISUBSYSTEM-003 | IFC-HMISUBSYSTEM-002 | verifies | Audio alert onset latency verification | VER-REQ-036 measures the 50ms alert onset budget specified in IFC-REQ-029. | 2026-03-31T03:09:08.310Z
VER-COMMUNICATIONSUBSYSTEM-004 | IFC-COMMUNICATIONSUBSYSTEM-005 | verifies | BLE 5.2 connection parameter and latency verification | VER-REQ-035 validates the 7.5ms connection interval and 20ms wireless transport latency specified in IFC-REQ-028. | 2026-03-31T03:09:07.469Z
VER-COMMUNICATIONSUBSYSTEM-003 | IFC-COMMUNICATIONSUBSYSTEM-004 | verifies | Verification of CDC-ECM USB interface and LTE latency | VER-REQ-034 tests the specific USB enumeration, APN configuration, and 100ms RTT requirement defined in IFC-REQ-027. | 2026-03-31T03:09:06.699Z
SYS-REQ-010 | SUB-VEHICLEPLATFORM-004 | derives | Turning radius budget derives from accessibility footprint and doorway navigation requirement | SYS-REQ-010 specifies 700mm width and accessibility environments. SUB-REQ-049 translates this into a specific turning radius that enables navigation through 900mm accessible doorways with clearance. | 2026-03-31T03:08:53.611Z
SYS-REQ-008 | SUB-HMISUBSYSTEM-004 | derives | Display must show accuracy so attendant can monitor degradation threshold | SYS-REQ-008 triggers mode transition at 70% accuracy. The care attendant can only act on this if they can see the live accuracy value — hence SUB-REQ-047 mandates 2Hz display update of the accuracy metric. | 2026-03-31T03:08:52.784Z
STK-REQ-002 | SYS-REQ-020 | derives | Biocompatibility requirement derives from user protection stakeholder need | STK-REQ-002 requires protection from injury in all modes. EEG electrode skin contact with non-biocompatible materials causes long-term dermal injury to users who may lack sensation to self-report. SYS-REQ-020 is the specific system-level implementation of this protection for skin-contacting components. | 2026-03-31T03:08:51.931Z
SYS-REQ-013 | SUB-BCIPROCESSINGSUBSYSTEM-011 | derives | Electrode biocompatibility derived from IEC 60601-1-2 EMC/medical compliance requirement | SUB-REQ-044 decomposes SYS-REQ-013 regulatory compliance to the electrode interface specifically. IEC 60601-1 biocompatibility requirements for patient-applied parts are part of the same medical electrical equipment standard family as the EMC requirement in SYS-REQ-013. | 2026-03-31T02:40:53.216Z
VER-SAFETYSUBSYSTEM-013 | SUB-SAFETYSUBSYSTEM-008 | verifies | Physical and electrical independence of Safety Subsystem PCB | Inspection of assembly and schematic is the appropriate method for a structural/physical isolation requirement; testing would require deliberate hardware fault injection which risks damage. | 2026-03-31T02:38:29.261Z
VER-COMMUNICATIONSUBSYSTEM-002 | SUB-COMMUNICATIONSUBSYSTEM-003 | verifies | End-to-end BLE emergency alert timing verification | End-to-end test from trigger to gateway receipt is required because intermediate component tests alone cannot verify total chain latency. | 2026-03-31T02:38:27.802Z
VER-BCIPROCESSINGSUBSYSTEM-005 | SUB-BCIPROCESSINGSUBSYSTEM-010 | verifies | Verification of BCI watchdog halt and safe-stop assertion | Fault injection test covers both invalid command and watchdog timeout paths, providing SIL-3 adequate test coverage for autonomous BCI safety constraints. | 2026-03-31T02:38:11.335Z
VER-PERCEPTIONSUBSYSTEM-004 | SUB-PERCEPTIONSUBSYSTEM-005 | verifies | Verification of tilt hazard signal assertion timing | Tilt rig test with I2C bus monitoring confirms signal propagation to the Safety Monitor Processor within the 50ms budget. | 2026-03-31T02:38:09.944Z
VER-HMISUBSYSTEM-002 | SUB-HMISUBSYSTEM-003 | verifies | Verification of joystick authority transfer latency | This test directly measures the 100ms latency bound specified in SUB-REQ-034 using calibrated CAN bus timing. | 2026-03-31T02:38:08.479Z
SYS-REQ-004 | IFC-POWERSUBSYSTEM-003 | derives | BMS to MAP CAN interface derived from battery state monitoring for range management | IFC-REQ-017 defines the CAN interface from BMS to MAP for battery state telemetry. This interface exists because SYS-REQ-004 requires the MAP to manage power budgeting for 4-hour operation, which requires real-time BMS data. | 2026-03-31T02:37:39.295Z
SYS-REQ-008 | SUB-HMISUBSYSTEM-002 | derives | Status LED array mode indication derived from care attendant alert requirement | SUB-REQ-033 decomposes SYS-REQ-008 care attendant alert requirement to the HMI Status LED Array. Distinct colour states for each operating mode enable rapid visual assessment of system state by the attendant from a distance. | 2026-03-31T02:37:38.479Z
SYS-REQ-004 | SUB-VEHICLEPLATFORM-002 | derives | Electronics bay thermal limit derived from 4-hour continuous operation requirement | SUB-REQ-028 decomposes SYS-REQ-004 continuous operation requirement to the Electronics Bay thermal constraint. The 70°C limit is derived from the operating temperature range of the Main Application Processor and BMS components housed within. | 2026-03-31T02:37:37.652Z
SYS-REQ-006 | SUB-PERCEPTIONSUBSYSTEM-003 | derives | Side proximity sensor detection zone derived from lateral obstacle detection requirement | SUB-REQ-025 decomposes SYS-REQ-006 1m lateral detection requirement to the Side Proximity Sensor Pair component. The 0.5m per side corresponds to the 1m total lateral zone, and zero false-negatives aligns with the system-level safety requirement. | 2026-03-31T02:37:36.844Z
SYS-REQ-003 | SUB-DRIVESUBSYSTEM-005 | derives | Drive motor torque capacity derived from speed and payload requirement | SUB-REQ-022 decomposes SYS-REQ-003 speed limit and SYS-REQ-010 payload requirements to the motor assemblies. Continuous torque capacity of each motor must be sufficient to maintain 6 km/h with 120 kg payload plus vehicle mass. | 2026-03-31T02:37:23.457Z
SYS-REQ-007 | SUB-BCIPROCESSINGSUBSYSTEM-006 | derives | Artifact engine watchdog interface derived from BCI degraded mode safety requirement | SUB-REQ-017 decomposes SYS-REQ-007 degraded mode transition requirement. The watchdog interface enables the BCI subsystem supervisor to detect artifact engine failure and initiate degraded mode transition without relying on software polling. | 2026-03-31T02:37:22.655Z
SYS-REQ-014 | SUB-POWERSUBSYSTEM-004 | derives | Charge controller timing derived from system charge time requirement | SUB-REQ-016 decomposes SYS-REQ-014 4-hour charge requirement to the Charge Controller component, specifying the 3-hour target (with margin) and CC/CV profile. | 2026-03-31T02:37:21.867Z
SYS-REQ-004 | SUB-POWERSUBSYSTEM-003 | derives | Power rail stability derived from 4-hour continuous operation requirement | SUB-REQ-015 decomposes SYS-REQ-004 continuous operation requirement to the DC-DC Converter Array. Stable regulated rails are a prerequisite for all electronics to achieve the 4-hour operation window. | 2026-03-31T02:37:21.032Z
SYS-REQ-001 | SUB-BCIPROCESSINGSUBSYSTEM-005 | derives | Per-user CSP filter loading derived from 85% classification accuracy requirement | SUB-REQ-012 decomposes SYS-REQ-001 85% accuracy requirement to the Feature Extraction Processor. Personalised CSP filters are the mechanism by which per-user classification accuracy is achieved and maintained. | 2026-03-31T02:37:20.186Z
SYS-REQ-002 | SUB-SAFETYSUBSYSTEM-008 | derives | Safety Subsystem physical separation derives from SIL-3 independence requirement | SUB-REQ-043 addresses the physical independence required for IEC 61508 SIL-3 independence between safety and non-safety processing channels, directly supporting the 200ms emergency stop guarantee in SYS-REQ-002. | 2026-03-31T02:37:05.447Z
SYS-REQ-002 | SUB-BCIPROCESSINGSUBSYSTEM-010 | derives | BCI watchdog safety constraint derived from 200ms emergency stop requirement | SUB-REQ-042 implements the autonomous operation safety constraint to ensure the 200ms emergency stop can be achieved when the BCI pipeline produces invalid commands. | 2026-03-31T02:37:04.611Z
SYS-REQ-002 | SUB-BCIPROCESSINGSUBSYSTEM-009 | derives | Artifact rejection engine power envelope derived from safety response reliability | SUB-REQ-041 bounds the ARE power draw to prevent BCI pipeline failure. The ARE must remain operational for SYS-REQ-002 200ms emergency stop to function correctly — ARE failure would prevent seizure detection from triggering emergency stop. | 2026-03-31T02:37:03.798Z
SYS-REQ-011 | SUB-PERCEPTIONSUBSYSTEM-006 | derives | Inclinometer power envelope derived from tilt detection reliability requirement | SUB-REQ-040 ensures the inclinometer sensor powering addresses the reliability implicit in SYS-REQ-011 tilt detection. Power supply margin prevents brown-out-induced false alarms. | 2026-03-31T02:37:03.001Z
SYS-REQ-019 | SUB-BCIPROCESSINGSUBSYSTEM-008 | derives | AES-256 encryption and RBAC decomposed to Main Application Processor | SUB-REQ-039 decomposes SYS-REQ-019 security requirements to the MAP as data custodian for all BCI pipeline data. | 2026-03-31T02:36:49.713Z
SYS-REQ-017 | SUB-COMMUNICATIONSUBSYSTEM-003 | derives | BLE emergency alert transmission decomposed to Communication Controller | SUB-REQ-038 decomposes SYS-REQ-017 500ms BLE alert transmission to the Communication Controller component that manages the BLE 5.0 module. | 2026-03-31T02:36:48.940Z
SYS-REQ-015 | SUB-BCIPROCESSINGSUBSYSTEM-007 | derives | USB-C service port diagnostic decomposed to Main Application Processor | SUB-REQ-037 decomposes SYS-REQ-015 USB-C diagnostic suite requirement to the component (MAP) that hosts the diagnostic executive and bus access. | 2026-03-31T02:36:48.131Z
SYS-REQ-014 | SUB-POWERSUBSYSTEM-005 | derives | Charge Controller behaviour decomposed from system charge time req | SUB-REQ-036 decomposes SYS-REQ-014 4-hour charge time requirement to the Power Subsystem Charge Controller, specifying the supply voltage acceptance and timing. | 2026-03-31T02:36:47.310Z
SYS-REQ-011 | SUB-PERCEPTIONSUBSYSTEM-005 | derives | Tilt detection sensing element behaviour for SYS-REQ-011 | SUB-REQ-035 decomposes SYS-REQ-011 tilt threshold to the Perception Subsystem sensing element. The timing budget allocation directly derives from this system requirement. | 2026-03-31T02:36:46.542Z
SYS-REQ-009 | SUB-HMISUBSYSTEM-003 | derives | Joystick handover latency decomposed to HMI/Drive interface | SUB-REQ-034 directly decomposes SYS-REQ-009 100ms authority transfer into an HMI subsystem behaviour requirement. SUB exists because of SYS-REQ-009. | 2026-03-31T02:36:45.719Z
SYS-REQ-009 | IFC-HMISUBSYSTEM-001 | derives | Caregiver override/status display drives Display Unit interface | SYS-REQ-009 requires carer display of override status; IFC-REQ-026 defines the display interface required to render that information. | 2026-03-31T02:18:36.546Z
SYS-REQ-001 | IFC-COMMUNICATIONSUBSYSTEM-003 | derives | BCI accuracy drives Communication Controller to BLE Module interface | SYS-REQ-001 BCI accuracy drives the data throughput of the IFC-REQ-025 UART interface between Communication Controller and BLE Module. | 2026-03-31T02:18:35.758Z
SYS-REQ-010 | IFC-VEHICLEPLATFORM-001 | derives | System dimensions requirement drives electronics bay mounting interface | SYS-REQ-010 defines physical envelope that constrains the Electronics Bay mounting interface specified in IFC-REQ-024. | 2026-03-31T02:18:35.000Z
SYS-REQ-006 | IFC-PERCEPTIONSUBSYSTEM-002 | derives | System obstacle detection drives lateral sensor interface spec | SYS-REQ-006 obstacle detection requirement drives the side proximity sensor interface timing in IFC-REQ-022. | 2026-03-31T02:18:34.238Z
VER-HMISUBSYSTEM-001 | SUB-HMISUBSYSTEM-001 | verifies | Audio alert SPL test verifies E-stop audible alert requirement | Sound level meter test validates SPL and onset time per IEC 60601-1-8. | 2026-03-31T02:17:51.624Z
VER-COMMUNICATIONSUBSYSTEM-001 | SUB-COMMUNICATIONSUBSYSTEM-002 | verifies | Firewall injection test verifies CAN bus isolation | Active penetration test validates the communication isolation specified in SUB-REQ-031. | 2026-03-31T02:16:15.096Z
SYS-REQ-001 | SUB-COMMUNICATIONSUBSYSTEM-001 | derives | BCI accuracy system req drives BLE throughput requirement | SYS-REQ-001 requires 85% BCI classification accuracy; SUB-REQ-030 ensures the BLE link delivers the continuous 8kHz EEG stream that the BCI classifier needs. | 2026-03-31T02:16:14.357Z
VER-VEHICLEPLATFORM-001 | SUB-VEHICLEPLATFORM-001 | verifies | ISO 7176-8 structural test verifies chassis load requirement | ISO 7176-8 test procedure directly validates the static and dynamic load thresholds specified in SUB-REQ-027. | 2026-03-31T02:13:06.448Z
SYS-REQ-010 | SUB-VEHICLEPLATFORM-003 | derives | System dimension requirement drives seat occupant accommodation range | SYS-REQ-010 sets vehicle width to 700mm; this constrains the seat width and drives the occupant accommodation range in SUB-REQ-029. | 2026-03-31T02:12:48.008Z
SYS-REQ-010 | SUB-VEHICLEPLATFORM-001 | derives | System dimensions requirement drives chassis load bearing specification | SYS-REQ-010 sets the physical envelope; SUB-REQ-027 specifies the structural performance the Chassis Frame must achieve within that envelope. | 2026-03-31T02:12:47.167Z
VER-PERCEPTIONSUBSYSTEM-003 | SUB-PERCEPTIONSUBSYSTEM-004 | verifies | Verification of MCU timeout triggering E-stop sequence | MCU power-off test validates the safe state transition specified in SUB-REQ-026. | 2026-03-31T02:09:51.158Z
VER-PERCEPTIONSUBSYSTEM-002 | IFC-PERCEPTIONSUBSYSTEM-003 | verifies | Verification of SPI protocol between Perception MCU and Safety Monitor Processor | Logic analyser capture validates the 4-byte frame format, 10Hz rate, and fault detection specified in IFC-REQ-023. | 2026-03-31T02:09:50.365Z
VER-PERCEPTIONSUBSYSTEM-001 | SUB-PERCEPTIONSUBSYSTEM-001 | verifies | Verification of forward sensor detection accuracy and update rate | Physical test with calibrated targets validates both the sensor accuracy and 10Hz rate specified in SUB-REQ-023. | 2026-03-31T02:09:49.524Z
SYS-REQ-006 | IFC-PERCEPTIONSUBSYSTEM-003 | derives | System obstacle response drives MCU-to-SMP interface timing | The system-level obstacle detection requirement drives the 10Hz frame rate and frame counter at the MCU-SMP interface to ensure continuous safety monitoring. | 2026-03-31T02:09:09.966Z
SYS-REQ-006 | IFC-PERCEPTIONSUBSYSTEM-001 | derives | System obstacle detection zone drives sensor bus specification | The 2m detection zone and 10Hz update rate in SYS-REQ-006 determine the I2C bandwidth and transfer frequency specified in IFC-REQ-021. | 2026-03-31T02:09:09.086Z
SYS-REQ-002 | SUB-PERCEPTIONSUBSYSTEM-004 | derives | Perception safe state derives from system emergency stop requirement | SYS-REQ-002 requires 20ms system-level E-stop; SUB-REQ-026 ensures Perception MCU communication loss triggers the E-stop chain rather than allowing silent failure. | 2026-03-31T02:08:40.043Z
SYS-REQ-006 | SUB-PERCEPTIONSUBSYSTEM-002 | derives | MCU latency derives from system-level obstacle response time | SYS-REQ-006 implicitly requires end-to-end detection-to-alert latency; SUB-REQ-024 allocates the Perception MCU's processing budget within that chain. | 2026-03-31T02:08:39.158Z
SYS-REQ-006 | SUB-PERCEPTIONSUBSYSTEM-001 | derives | Forward sensor array specs derive from system obstacle detection zone | SYS-REQ-006 specifies 2m forward detection zone; SUB-REQ-023 specifies the sensor hardware performance that implements this zone. | 2026-03-31T02:08:38.379Z
SYS-REQ-001 | IFC-BCIPROCESSINGSUBSYSTEM-005 | derives | Classifier-Command Arbitration interface derives from BCI latency and accuracy | SYS-REQ-001 accuracy and the 50ms end-to-end latency budget together constrain the command probability format and rate at IFC-REQ-014. High-probability thresholds on this interface enforce the classification confidence required by SYS-REQ-001. | 2026-03-31T01:41:55.887Z
SYS-REQ-001 | IFC-BCIPROCESSINGSUBSYSTEM-004 | derives | Feature Extraction-Classifier interface derives from BCI accuracy requirement | SYS-REQ-001 accuracy requirement cascades to feature vector dimensionality at IFC-REQ-013. Insufficient feature dimensionality or update rate directly limits classification accuracy. | 2026-03-31T01:41:55.084Z
SYS-REQ-001 | IFC-BCIPROCESSINGSUBSYSTEM-003 | derives | Artifact Rejection-Feature Extraction interface derives from BCI classification pipeline | SYS-REQ-001 classification accuracy target drives the clean EEG data throughput requirement at IFC-REQ-012. Cleaned signal at this interface directly conditions classification quality. | 2026-03-31T01:41:54.284Z
SYS-REQ-001 | IFC-BCIPROCESSINGSUBSYSTEM-002 | derives | EEG Acquisition-Artifact Rejection interface derives from BCI classification performance | SYS-REQ-001 requires 80% BCI classification accuracy. The 32-channel EEG data rate on IFC-REQ-011 (250Hz, 16-bit) is the minimum needed to support the feature extraction and classification pipeline that achieves this accuracy. | 2026-03-31T01:41:53.440Z
SYS-REQ-002 | SUB-DRIVESUBSYSTEM-004 | derives | Drive Power Stage overcurrent protection derives from system safe-stop requirement | SYS-REQ-002 requires 250ms safe stop on any trigger. Hardware overcurrent protection in Drive Power Stage is the last-resort safe-stop mechanism independent of firmware — required to satisfy SIL 2 safe-state reachability. This is a direct derivation of the system-level safety trigger requirement. | 2026-03-31T01:41:52.547Z
VER-DRIVESUBSYSTEM-004 | SUB-DRIVESUBSYSTEM-003 | verifies | VER-022 verifies MCU 250ms safe-state on CAN loss | End-to-end safe-state timing must be verified on complete vehicle hardware at SIL 2 — simulation is insufficient for safety case credit. | 2026-03-31T01:40:38.258Z
VER-DRIVESUBSYSTEM-003 | IFC-DRIVESUBSYSTEM-003 | verifies | VER-021 verifies Drive Power Stage overcurrent trip interface | Overcurrent trip latency is hardware-testable at bench, verifying the <5ms requirement without vehicle assembly. | 2026-03-31T01:40:37.473Z
VER-DRIVESUBSYSTEM-002 | IFC-DRIVESUBSYSTEM-002 | verifies | VER-020 verifies right motor PWM/encoder interface symmetry | Direct verification of IFC-REQ-019. Adds left-right latency asymmetry check specific to differential-drive correctness. | 2026-03-31T01:40:36.689Z
VER-DRIVESUBSYSTEM-001 | IFC-DRIVESUBSYSTEM-001 | verifies | VER-019 verifies left motor PWM/encoder interface | Direct verification of IFC-REQ-018 pass/fail criteria: PWM frequency, encoder timing, and velocity accuracy measured on hardware bench. | 2026-03-31T01:40:35.851Z
SYS-REQ-002 | SUB-DRIVESUBSYSTEM-003 | derives | MCU CAN loss safe-state derives from system 250ms safe stop | SYS-REQ-002 requires 250ms safe stop on any safety trigger. SUB-REQ-020 is the Drive Subsystem's implementation of that requirement — the CAN heartbeat loss trigger and regenerative braking response is the mechanism by which the system meets the 250ms budget. | 2026-03-31T01:40:24.735Z
SYS-REQ-003 | SUB-DRIVESUBSYSTEM-002 | derives | MCU speed clamping directly implements SYS speed limit | SUB-REQ-019 specifies firmware-level speed clamping in the Motor Controller Unit. This requirement derives directly from SYS-REQ-003 — the system cannot comply with speed limits without active enforcement in the actuator controller. | 2026-03-31T01:40:23.929Z
SYS-REQ-003 | SUB-DRIVESUBSYSTEM-001 | derives | MCU velocity control accuracy derives from system speed limit | SYS-REQ-003 mandates 6km/h and 2km/h speed limits. SUB-REQ-018 implements the closed-loop control that enforces these limits with ±0.1km/h accuracy. The subsystem requirement exists because the system-level speed limit requires a mechanism to measure and enforce actual wheel velocity. | 2026-03-31T01:40:23.168Z
VER-POWERSUBSYSTEM-002 | IFC-POWERSUBSYSTEM-002 | verifies | GPIO fault line test verifies BMS-Safety interface independence | VER-REQ-018 verifies the critical safety property of CAN-independence for the thermal fault signal. | 2026-03-31T01:09:48.335Z
VER-POWERSUBSYSTEM-001 | SUB-POWERSUBSYSTEM-001 | verifies | Thermal cutoff hardware test verifies BMS protection response | VER-REQ-017 provides direct hardware measurement of the BMS protection cutoff timing. | 2026-03-31T01:09:47.571Z
SYS-REQ-004 | SUB-POWERSUBSYSTEM-002 | derives | Battery runtime derives from system 4-hour endurance requirement | SYS-REQ-004 specifies 4-hour operation; SUB-REQ-014 implements this as a battery pack capacity requirement with the 10% SoC floor. | 2026-03-31T01:09:46.798Z
SYS-REQ-012 | SUB-POWERSUBSYSTEM-001 | derives | BMS thermal cutoff derives from system battery safety requirement | SYS-REQ-012 requires pack disconnect on thermal fault; SUB-REQ-013 implements the specific 200ms response time for the BMS protection circuit. | 2026-03-31T01:09:46.004Z
VER-BCIPROCESSINGSUBSYSTEM-004 | IFC-BCIPROCESSINGSUBSYSTEM-006 | verifies | CAN bus analysis verifies Drive Subsystem command interface | VER-REQ-016 uses CAN analyser to verify all protocol, timing, and counter requirements in IFC-REQ-015. | 2026-03-31T01:06:52.908Z
VER-BCIPROCESSINGSUBSYSTEM-003 | SUB-BCIPROCESSINGSUBSYSTEM-003 | verifies | Pipeline latency measurement verifies 150ms end-to-end constraint | VER-REQ-015 provides timestamped measurement from acquisition input to CAN output, directly verifying SUB-REQ-010's 150ms latency. | 2026-03-31T01:06:52.122Z
VER-BCIPROCESSINGSUBSYSTEM-002 | SUB-BCIPROCESSINGSUBSYSTEM-004 | verifies | SNR dropout test verifies signal-loss safe state | VER-REQ-014 provides direct empirical evidence of SUB-REQ-011 compliance via CAN bus timing measurement. | 2026-03-31T01:06:51.269Z
SYS-REQ-001 | SUB-BCIPROCESSINGSUBSYSTEM-001 | derives | Artifact rejection derives from system EEG classification requirement | SYS-REQ-001 classification accuracy depends on clean EEG input; SUB-REQ-008 specifies the artifact rejection performance needed to maintain that input quality. | 2026-03-31T01:06:50.484Z
SYS-REQ-007 | SUB-BCIPROCESSINGSUBSYSTEM-004 | derives | Signal-loss safe-stop derives from system SNR monitoring requirement | SYS-REQ-007 requires safe-stop on sustained BCI signal loss; SUB-REQ-011 implements this at Command Arbitration Module with the 3-second timeout and STOP command output. | 2026-03-31T01:06:49.694Z
SYS-REQ-001 | SUB-BCIPROCESSINGSUBSYSTEM-003 | derives | End-to-end latency requirement derives from system response time | SYS-REQ-001 requires motor imagery classification within 150ms; SUB-REQ-010 implements this at the Command Arbitration Module output, the last stage of the BCI pipeline. | 2026-03-31T01:06:48.908Z
SYS-REQ-001 | SUB-BCIPROCESSINGSUBSYSTEM-002 | derives | BCI classification accuracy floor derives from system command reliability | SYS-REQ-001 requires reliable EEG command classification; SUB-REQ-009 implements the 75% accuracy floor that enables this capability in the BCI Classifier component. | 2026-03-31T01:06:48.031Z
STK-REQ-012 | SYS-REQ-019 | derives | CMMS integration need → data security requirement for CMMS telemetry | STK-REQ-012 requires CMMS REST API integration for maintenance scheduling and fault reporting; SYS-REQ-019's TLS 1.3 and RBAC requirements are directly derived from the need to securely transmit maintenance telemetry to the facility CMMS without exposing patient EEG data. | 2026-03-31T00:38:02.768Z
STK-REQ-010 | SYS-REQ-013 | derives | Regulatory design history file need → EMC/medical device compliance requirement | STK-REQ-010 requires IEC 60601-1 compliance documentation; SYS-REQ-013 specifies IEC 60601-1-2 EMC compliance as a concrete system-level requirement that is part of the technical file required by MDR 2017/745 for CE marking, directly supporting the design history file obligation. | 2026-03-31T00:37:51.515Z
SYS-REQ-019 | IFC-COMMUNICATIONSUBSYSTEM-002 | derives | Data encryption/TLS requirement → CMMS REST API interface | SYS-REQ-019 requires TLS 1.3 in transit for all data interfaces; IFC-REQ-006 specifies the CMMS REST API over Wi-Fi with TLS 1.3 — the interface directly implements the encryption mandate for facility telemetry data. | 2026-03-31T00:37:43.921Z
SYS-REQ-015 | IFC-REQ-005 | derives | USB-C diagnostic test suite requirement → service diagnostic USB-C interface | SYS-REQ-015 requires an automated diagnostic test suite accessible via USB-C service port; IFC-REQ-005 specifies the USB-C wired interface protocol, bandwidth, and access-level authentication that implements this diagnostic access point. | 2026-03-31T00:37:43.171Z
SYS-REQ-019 | IFC-COMMUNICATIONSUBSYSTEM-001 | derives | Data encryption/RBAC requirement → smartphone companion app BLE interface | SYS-REQ-019 requires TLS 1.3 in transit and RBAC for clinical/maintenance interfaces; IFC-REQ-004 specifies the BLE 5.0 companion app interface that carries session status and EEG data requiring the TLS and RBAC protections mandated by SYS-REQ-019. | 2026-03-31T00:37:42.318Z
STK-REQ-002 | SYS-REQ-014 | derives | User safety → charging interface safety requirement | STK-REQ-002 requires user safety across all modes; SYS-REQ-014's galvanic isolation requirement implements electrical safety during the charging mode — isolation prevents shock risk to a user who remains in the buggy during charge. | 2026-03-31T00:37:22.686Z
SYS-REQ-014 | IFC-POWERSUBSYSTEM-001 | derives | 4-hour charging requirement → galvanically isolated charging interface | SYS-REQ-014 specifies charging from 20% to 100% within 4h via 230V facility supply; IFC-REQ-002 defines the IEC 60320 C13/C14 galvanic isolation interface that implements safe electrical coupling to the facility charging infrastructure. | 2026-03-31T00:37:21.876Z
VER-SAFETYSUBSYSTEM-012 | IFC-SAFETYSUBSYSTEM-004 | verifies | SPI protocol capture + CRC injection → inclinometer SPI interface requirement | VER-REQ-013 verifies frame format, CRC integrity, and fault-detection response specified in IFC-REQ-009. | 2026-03-31T00:36:19.305Z
VER-SAFETYSUBSYSTEM-011 | IFC-SAFETYSUBSYSTEM-001 | verifies | Emergency BLE alert latency test → emergency/nurse call interface requirement | VER-REQ-012 tests the < 500ms E-stop-to-advertisement latency and payload content correctness specified in IFC-REQ-003. | 2026-03-31T00:36:18.505Z
VER-BCIPROCESSINGSUBSYSTEM-001 | IFC-BCIPROCESSINGSUBSYSTEM-001 | verifies | BLE EEG streaming test → BLE EEG headset interface requirement | VER-REQ-011 tests all acceptance criteria in IFC-REQ-001: < 0.1% packet loss at 5m, < 1% at 10m, < 3s connection time. | 2026-03-31T00:36:17.627Z
VER-SAFETYSUBSYSTEM-010 | SUB-SAFETYSUBSYSTEM-007 | verifies | Boundary condition watchdog test → Safety Monitor Processor heartbeat requirement | VER-REQ-010 verifies both threshold conditions in SUB-REQ-007: 500ms absence and ±20% period deviation, with 50-run repeatability. | 2026-03-31T00:35:35.880Z
VER-SAFETYSUBSYSTEM-009 | SUB-SAFETYSUBSYSTEM-006 | verifies | Fault injection test → Safety Subsystem E-stop state machine requirement | VER-REQ-009 covers all five trigger conditions in SUB-REQ-006 individually and in combination, measuring the full 200ms response chain. | 2026-03-31T00:35:35.074Z
VER-SAFETYSUBSYSTEM-008 | SUB-SAFETYSUBSYSTEM-005 | verifies | Oscilloscope + schematic inspection → manual E-stop hardwire requirement | VER-REQ-008 verifies both measurable criteria in SUB-REQ-005: ≤5ms disconnect time and absence of software mediation in the circuit. | 2026-03-31T00:35:34.249Z
VER-SAFETYSUBSYSTEM-007 | SUB-SAFETYSUBSYSTEM-004 | verifies | Tilt table test → Inclinometer tilt sensor unit requirement | VER-REQ-007 directly tests all measurable acceptance criteria in SUB-REQ-004: ±0.5° accuracy, 100Hz rate, and 200ms E-stop trigger chain. | 2026-03-31T00:35:33.434Z
STK-REQ-015 | SYS-REQ-003 | derives | Audible travel tone/direction indicator need → speed limit system requirement | STK-REQ-015 requires audible tones and visible direction indicators at low speed — SYS-REQ-003's 6 km/h maximum speed limit defines the operational envelope within which these pedestrian proximity warnings must operate and be audible. | 2026-03-31T00:34:44.592Z
STK-REQ-011 | SYS-REQ-015 | derives | Post-market surveillance data collection need → automated diagnostic test suite | STK-REQ-011 requires automated incident logging for post-market surveillance; SYS-REQ-015's automated diagnostic test suite with USB-C access provides the data capture mechanism that feeds post-market surveillance reports as required by MDR 2017/745. | 2026-03-31T00:34:43.810Z
STK-REQ-003 | SYS-REQ-007 | derives | Cognitive performance detection need → BCI SNR threshold degradation response | STK-REQ-003 requires detection of declining cognitive performance before accuracy drops below safe thresholds; SYS-REQ-007 specifies the concrete system response when BCI SNR drops below classification threshold for >5 seconds — the quantified realisation of this stakeholder safety need. | 2026-03-31T00:34:43.076Z
STK-REQ-008 | SYS-REQ-018 | derives | Clinical BCI parameter configuration need → startup calibration requirement | STK-REQ-008 requires clinical prescribers to configure patient-specific BCI parameters; SYS-REQ-018 specifies the combined startup/calibration sequence within 3 minutes that is the operational realisation of patient-specific BCI profile loading. | 2026-03-31T00:34:42.283Z
STK-REQ-017 | SYS-REQ-004 | derives | Outdoor operation environmental need → battery endurance/range requirement | STK-REQ-017 requires outdoor paved path operation at -5°C to 40°C — SYS-REQ-004's 4-hour/15 km endurance requirement must be met across this temperature range, which directly derives from the outdoor operational context. | 2026-03-31T00:34:25.182Z
STK-REQ-014 | SYS-REQ-003 | derives | Safe stopping distance need → speed limit system requirement | STK-REQ-014 requires safe stopping distances and predictable movement; SYS-REQ-003 specifies the 6 km/h Normal and 2 km/h Degraded speed limits that bound stopping distances within safe margins for indoor care facility environments. | 2026-03-31T00:34:24.383Z
STK-REQ-009 | SYS-REQ-019 | derives | Per-session usage data recording need → EEG data encryption and GDPR compliance | STK-REQ-009 requires session data recording including EEG classification accuracy — SYS-REQ-019 specifies the AES-256 and TLS 1.3 encryption required to store/transmit that biometric data in compliance with GDPR and NHS Digital standards. | 2026-03-31T00:34:23.540Z
STK-REQ-007 | SYS-REQ-016 | derives | OTA/wired firmware update need → dual-bank flash firmware update system req | STK-REQ-007 requires OTA and wired firmware update with automatic rollback; SYS-REQ-016 specifies the dual-bank flash architecture with rollback that implements this firmware management need. | 2026-03-31T00:34:22.734Z
STK-REQ-006 | SYS-REQ-015 | derives | Maintenance diagnostic interface need → automated diagnostic test suite | STK-REQ-006 requires a wired diagnostic interface for maintenance technicians; SYS-REQ-015 specifies the USB-C accessible diagnostic test suite that fulfils this maintenance access requirement. | 2026-03-31T00:34:21.885Z
STK-REQ-005 | SYS-REQ-017 | derives | Attendant alert stakeholder need → BLE emergency alert system requirement | STK-REQ-005 requires distinct audible/visual signals to the care attendant; SYS-REQ-017 specifies the BLE 5.0 iBeacon-compatible emergency transmission that implements this attendant alerting function. | 2026-03-31T00:34:21.065Z
SYS-REQ-005 | IFC-SAFETYSUBSYSTEM-005 | derives | Seizure detection system req → SDM-SMP shared memory interface | IFC-REQ-010 exists because SYS-REQ-005 requires seizure detection to feed into the safety chain — the interface between the detection module and the safety processor implements this coupling. | 2026-03-31T00:09:29.185Z
SYS-REQ-011 | IFC-SAFETYSUBSYSTEM-004 | derives | Tilt detection requirement → inclinometer SPI interface spec | IFC-REQ-009 exists because SYS-REQ-011 requires the inclinometer to be connected directly to the safety path processor — this interface definition specifies how that connection is implemented. | 2026-03-31T00:09:28.312Z
VER-SAFETYSUBSYSTEM-006 | SYS-REQ-002 | verifies | End-to-end vehicle stop test verifies SYS E-stop requirement | VER-REQ-006 verifies SYS-REQ-002 at system integration level by testing all 5 safety trigger paths. | 2026-03-31T00:08:44.077Z
VER-SAFETYSUBSYSTEM-005 | IFC-SAFETYSUBSYSTEM-003 | verifies | Single-channel fault test verifies dual-channel relay interface | VER-REQ-005 verifies IFC-REQ-008 by testing each channel independently for fail-safe relay de-energisation. | 2026-03-31T00:08:43.291Z
VER-SAFETYSUBSYSTEM-004 | IFC-SAFETYSUBSYSTEM-002 | verifies | Heartbeat halt test verifies heartbeat interface spec | VER-REQ-004 verifies IFC-REQ-007 by testing both the 500ms watchdog timing and the no-data-payload constraint on the heartbeat line. | 2026-03-31T00:08:42.513Z
VER-SAFETYSUBSYSTEM-003 | SUB-SAFETYSUBSYSTEM-003 | verifies | CHB-MIT replay verifies seizure detection spec | VER-REQ-003 verifies SUB-REQ-003 using published benchmark dataset for reproducible sensitivity and false positive rate measurement. | 2026-03-31T00:08:41.717Z
VER-SAFETYSUBSYSTEM-002 | SUB-SAFETYSUBSYSTEM-002 | verifies | Load test verifies relay 20ms disconnect | VER-REQ-002 verifies SUB-REQ-002 by measuring relay switching time under full rated current range. | 2026-03-31T00:08:40.614Z
VER-SAFETYSUBSYSTEM-001 | SUB-SAFETYSUBSYSTEM-001 | verifies | Fault injection test verifies SMP independence | VER-REQ-001 verifies SUB-REQ-001 by physically inducing main processor failure and confirming safety processor continues independently. | 2026-03-31T00:08:39.787Z
SYS-REQ-002 | SUB-SAFETYSUBSYSTEM-005 | derives | SYS manual override requirement → hardwired E-stop architecture | SUB-REQ-005 exists because SYS-REQ-002 requires the manual E-stop to function as a safety trigger — hardwired architecture is the only design that guarantees this independently of software state. | 2026-03-31T00:07:19.762Z
SYS-REQ-011 | SUB-SAFETYSUBSYSTEM-004 | derives | SYS tilt threshold → Inclinometer sensor specification | SUB-REQ-004 decomposes SYS-REQ-011 into the sensing hardware requirements: measurement rate, accuracy, and debounce timing for the Inclinometer Tilt Sensor Unit. | 2026-03-31T00:07:18.993Z
SYS-REQ-005 | SUB-SAFETYSUBSYSTEM-003 | derives | SYS seizure detection trigger → Seizure Detection Module spec | SUB-REQ-003 decomposes SYS-REQ-005 into the specific algorithm, detection threshold, latency, and false-positive rate requirements for the Seizure Detection Module component. | 2026-03-31T00:07:18.191Z
SYS-REQ-002 | SUB-SAFETYSUBSYSTEM-002 | derives | SYS E-stop 200ms budget → relay 20ms disconnect | SUB-REQ-002 derives from SYS-REQ-002 200ms total response budget — the relay is the actuator that physically implements power disconnection, and its latency is the largest single time component. | 2026-03-31T00:07:17.331Z
SYS-REQ-002 | SUB-SAFETYSUBSYSTEM-001 | derives | SYS E-stop response requirement → Safety Monitor Processor independence | SUB-REQ-001 exists because SYS-REQ-002 requires emergency stop to be achievable even during main processor failure — this is only possible if the safety processor is architecturally independent. | 2026-03-31T00:07:16.557Z
STK-REQ-018 | SYS-REQ-019 | derives | GDPR drives encryption and access control | EEG biometric data is GDPR Article 9 special category data. AES-256 and TLS 1.3 are minimum standards for health data protection. | 2026-03-30T23:38:16.961Z
STK-REQ-016 | SYS-REQ-013 | derives | EMI environment drives EMC compliance | Hospital operating environment contains MRI fringe fields and electrosurgical units. IEC 60601-1-2 EMC compliance ensures EEG acquisition degrades safely rather than producing undetectable misclassifications. | 2026-03-30T23:38:16.131Z
STK-REQ-002 | SYS-REQ-012 | derives | User safety during battery thermal event | H-003 (thermal runaway) is catastrophic/SIL 3. Non-ambulatory user cannot evacuate — 500ms disconnect, fire suppression, and facility alert are all essential mitigation layers. | 2026-03-30T23:38:15.316Z
STK-REQ-013 | SYS-REQ-010 | derives | Facility compatibility drives vehicle dimensions | 700mm width fits 900mm DDA doorways with clearance. 1200mm length fits standard lift cabs. 1500mm turning radius enables corridor U-turns. Dimensions derived from DDA and Building Regulations standards. | 2026-03-30T23:38:14.506Z
STK-REQ-004 | SYS-REQ-009 | derives | Carer override requires rapid handover | Care attendant needs immediate control during emergencies. 100ms handover ensures seamless transition. Obstacle detection stays active because carer may be unfamiliar with environment. | 2026-03-30T23:38:13.579Z
STK-REQ-003 | SYS-REQ-008 | derives | Fatigue detection triggers degraded mode | Cognitive fatigue causes classification accuracy to drift below safe thresholds over 25-30 minutes. The 70% threshold with 2-minute rolling window is derived from the Cognitive Fatigue scenario. | 2026-03-30T23:37:52.370Z
STK-REQ-002 | SYS-REQ-006 | derives | User safety requires collision prevention | H-002 (collision) is SIL 2. BCI-controlled vehicle cannot rely on user visual scanning alone — autonomous obstacle detection is the primary collision prevention mechanism. | 2026-03-30T23:37:51.513Z
STK-REQ-002 | SYS-REQ-005 | derives | User safety during seizure events | H-006 (seizure) is catastrophic/SIL 3. Seizure EEG can be misclassified as movement commands — detection and E-stop within 150ms prevents erratic vehicle behaviour with a seizing user. | 2026-03-30T23:37:50.700Z
STK-REQ-002 | SYS-REQ-002 | derives | User safety requires rapid emergency stop | Non-ambulatory user cannot self-protect — the system must halt within 200ms to prevent injury from signal loss, collision, or fault conditions. | 2026-03-30T23:37:49.894Z
STK-REQ-001 | SYS-REQ-001 | derives | Independent navigation requires reliable BCI classification | The user cannot navigate independently unless the BCI accurately classifies motor imagery into movement commands. 85% accuracy is the minimum for safe 4-class navigation. | 2026-03-30T23:37:49.001Z

## Linksets
Source Document | Target Document | Default Link Type | Link Count
subsystem-requirements | architecture-decisions |  | 0
system-requirements | architecture-decisions |  | 7
subsystem-requirements | verification-requirements |  | 79
interface-requirements | verification-requirements |  | 31
system-requirements | verification-requirements |  | 6
system-requirements | subsystem-requirements |  | 77
system-requirements | interface-requirements |  | 19
stakeholder-requirements | system-requirements |  | 29

## Sections
Document | Section | Short Code | Order | Requirements
architecture-decisions | Architecture Decisions | ARC | 0 | 7
interface-requirements | Interface Requirements | IFC | 0 | 1
interface-requirements | BCI Processing Subsystem |  | 0 | 6
interface-requirements | Communication Subsystem |  | 0 | 5
interface-requirements | Drive Subsystem |  | 0 | 3
interface-requirements | HMI Subsystem |  | 0 | 3
interface-requirements | Perception Subsystem |  | 0 | 3
interface-requirements | Power Subsystem |  | 0 | 3
interface-requirements | Safety Subsystem |  | 0 | 5
interface-requirements | Vehicle Platform |  | 0 | 2
stakeholder-requirements | Stakeholder Requirements | STK | 0 | 18
subsystem-requirements | Subsystem Requirements | SUB | 0 | 0
subsystem-requirements | BCI Processing Subsystem |  | 0 | 13
subsystem-requirements | Communication Subsystem |  | 0 | 7
subsystem-requirements | Drive Subsystem |  | 0 | 10
subsystem-requirements | HMI Subsystem |  | 0 | 11
subsystem-requirements | Perception Subsystem |  | 0 | 12
subsystem-requirements | Power Subsystem |  | 0 | 15
subsystem-requirements | Safety Subsystem |  | 0 | 15
subsystem-requirements | Vehicle Platform |  | 0 | 10
system-requirements | System Requirements | SYS | 0 | 23
verification-requirements | Verification Requirements | VER | 0 | 4
verification-requirements | BCI Processing Subsystem |  | 0 | 17
verification-requirements | Communication Subsystem |  | 0 | 13
verification-requirements | Drive Subsystem |  | 0 | 13
verification-requirements | HMI Subsystem |  | 0 | 14
verification-requirements | Perception Subsystem |  | 0 | 10
verification-requirements | Power Subsystem |  | 0 | 10
verification-requirements | Safety Subsystem |  | 0 | 23
verification-requirements | Vehicle Platform |  | 0 | 12